## df_osf_sia_wd
device_id | Manufacturer | Device | Website | Release date | Market status | Main use | Device costs | Wearable type | Location | Weight | Size | PPG | ECG | ICG | EMG | Respiration | EDA | EEG | BP | Accelerometer | Gyroscope | GPS | Skin temperature | Other signals | Water resistance | Battery life | Charging method | Charging duration | Bio-cueing | Bio-feedback | Raw data available | Provided parameters | Parameter sampling window (1min; 5min; 1hour; 1day) | Data transfer method | Compatibility | Required software | Additional software | Internal storage method | Device storage capacity | Server Data Storage | GDPR compliance | FDA approval/clearance | CE approval/label | Highest level of Validation Evidence | Number of validity and reliability studies reviewed | Studied parameters | General validity and reliability synthesis | Number of usability studies reviewed | General usability synthesis | Hyperlink to the device VRU page | Most recent date of RVU search | SiA Expert score (short-term) | SiA Expert score (long-term)
001_sia_wd_emp | Empatica | E4 | https://www.empatica.com/research/e4/ | 2015-01-01 | Discontinued | Research | 1542; one-time purchase price including a device + software | Watch | Wrist | 40 | L44xW40xH16 | 1; reflection; 64; wrist | 0 | 0 | 0 | 0 | 1; dry electrode; 4; wrist | 0 | 0 | 1; 3 axes; 32; wrist | 0 | 0 | 1; 0.2(36-39); 4; wrist | 0 | 1 | 48; hours | device charger | 120; minutes | 0 | 1 | 1 |  |  | BLE; USB 2.0 | Windows 7-10; iOS 8.2 or higher; Android 8.1 or higher; macOS NA | E4 manager; Empatica Connect web dashboard; | E4 Realtime | built-in storage | 60; hours; 60; mb | USA | 1 | 0 | 1 | external | 10 | HF; HR; IBI; ns.SCRs; NS-SCRs-AMP; NS-SCRs-FREQ; RMSSD; S; SCR-AMP; SCRs; SCL; SD1; SD1UP; SD1DOWN; SD2; SDNN ; SDRR | Most studies comparing E4 to a reference device (RD) suggest high validity of HR under static and dynamic conditions (one study found opposite results) and often report HRV parameters as valid but only in static conditions, although results seem conclusive only for RMSSD and SDNN. Movement and speech have been shown to result not only in decreased accuracy but also in high amounts of missing data. Missing data was also observed in static conditions and was attributed to issues with E4’s automatic IBI detection. It also appears E4 cannot provide convergently valid EDA (possibly due to differences in device location). However, one study reported high correlation of SCL with an RD despite it measuring from fingers, with yet another study reporting validity of amplitude of SCRs for strong sustained stressors. | 5 | E4 was generally considered easy to use with charging and handling the data reported as burdensome in multiple studies. Adherence was found to be high but results are inconsistent when it comes to wearing comfort. |  | 2023-06-27 | 5 | 6.7
002_sia_wd_emp | Empatica | EmbracePlus | https://www.empatica.com/embraceplus/ | 2020-01-01 | Current | Research & Clinical | 1984; three year bundle per a device + software (25% academic discount available) | Watch | Wrist | 38 | D32xH15 | 1; transmission & reflection; green, red and infrared; 26-208; wrist | 0 | 0 | 0 | 0 | 1; dry electrode; 1-4; wrist | 0 | 0 | 1; 3 axes; 26-208; wrist | 1; 3 axes; 26-208; wrist | 0 | 1; 0.1(30-45); 1-4; wrist | 0 | 1 | 48; hours | device charger | 90; minutes | 0 | 1 | 1 | HR; HRV; RR; Skin Temp; SpO2; EDA; Activity Counts; Sleep Detection | 1min: HR, HRV; 1day:Activity Counts | BLE; Blt 5.0 | Windows NA; iOS 11 or higher; Android 8.0 or higher; macOS NA | Care app; Care web portal |  | built-in storage | 24; hours; 128; mb |  | 1 | 1 | 1 | internal | 1 | SpO2 | Reliable and valid for use in clinical SpO2 measurement under static conditions, for both male and female subjects as well as for different skin tones. | 0 |  |  | 2023-09-18 | 5.7 | 8
003_sia_wd_vri | Vrije Universiteit van Amsterdam | VU-AMS 5fs | http://www.vu-ams.nl/ | 2007-01-01 | Current | Research | 4550 | CPU + electrodes | Hip | 260 | L126xW69xH32 | 0 | 1; spot electrodes; 2 channels; lead I; lead II; 1000 | 1; spot electrodes; 1000; thorax | 0 | 1; spot electrodes; impedence-derived; 250; thorax | 1; spot electrodes; 10; wrist-thenar | 0 | 0 | 1; 3 axes; 1000; hip | 0 | 0 | 0 | 0 | 0 | 24; hours | disposable batteries | 0; minutes | 0 | 0 | 1 | HR; HRV; RR; Skin Temp; SpO2; EDA; Activity Counts; Sleep Detection | 1min: CO, ECG-Q, HR, IBI, ICG-B, ICG-C, ICG-X, LVET, PEP, RMSSD, RR, RSA, SDNN, SV, TWA | SD-card; Blt | Windows NA; iOS NA; Android NA; macOS NA | VU-DAMS |  | SD-card | hours; 1024; mb |  | 1 | 0 | 1 | external | 3 | IBI; RBI; PEP; LVET; SV; CO; HR | Reliable and valid for measuring cardiac parameters using ICG and ECG. Absolute SV and CO values can deviate, especially during activity. Additionally, the VU-AMS 5fs device has been used as a gold standard reference device for various studies. | 0 | The validation studies showed it is feasible to measure with the VU-AMS device in daily life for 24h. |  | 2023-11-20 | 8.7 | 3
004_sia_wd_vri | Vrije Universiteit van Amsterdam | VU-AMS Core | http://www.vu-ams.nl/ | 2024-03-01 | Current | Research | 5995; one-time purchase price including a device + software | CPU + electrodes | Chest | 52 | L57xW57xH17 | 0 | 1; spot electrodes; 3 channels; lead I; lead II; aVF; 1000 | 1; spot electrodes; 1000; thorax | 0 | 1; spot electrodes; impedence-derived; 250; thorax | 0 | 0 | 0 | 1; 3 axes; 1000; upper sternum | 1; 3 axes; 1000; upper sternum | 0 | 0 | 0 | 0 | 28; hours | device charger | 180; minutes | 0 | 0 | 1 | CO; ECG-Q; HR; IBI; ICG-B; ICG-C; ICG-X; LVET; PEP; RMSSD; RR; RSA; SDNN ; SV; TWA | 1min: all parameters | mircoSD-card | Windows NA; iOS NA; Android NA; macOS NA | VU-DAMS |  | mircoSD-card | 160; hours; 32000; mb |  | 1 | 0 | 0 | No validation | 0 |  | No validation studies yet. This device is the successor of the VU-AMS 5fs device. | 0 |  |  | 2023-11-20 | 7.7 | 4.3
005_sia_wd_fit | Fitbit | Charge 5 | https://www.fitbit.com/global/nl/products/trackers/charge5?sku=421BKBK | 2021-09-01 | Current | Consumer | 150 | Watch | Wrist | 28 | L36.7xW22.7xH11.2 | 1; reflection; wrist | 1; Manually Initiated Recording; 1 channel; lead I | 0 | 0 | 0 | 1; dry electrode; only the scan function, no continuous EDA; wrist | 0 | 0 | 1; 3 axes; wrist | 0 | 0 | 1; wrist | 1; ambient light | 1 | 168; hours | device charger | 120; minutes | 0 | 1 | 0 | EDA-MIR; HR; HRV; RR; Skin Temp; SpO2; Cardio-Fitness Score; Hourly Activity; Sleep Duration; Sleep Score; Sleep Stages; Stress Management Score | 1hour: HR, Steps, Hourly Activity; 1day: Steps, Distance, Floors, Calories burned, Sleep Score; Stress Management Score; Sleep stages | Blt 5.0; NFC | Windows NA; iOS 15 or higher; Android 9.0 or higher; macOS NA | Fitbit app | Fitbit ECG app; Fitbit EDA Scan app | built-in storage; need to connect and send data to the Fitbit app to see the data | hours; mb | several countries within the European Economic Area, United Kindom and Switzerland, controlled by Fitbit International Limited, an Irisch company. | 1 | 1 | 1 | external | 1 | HR, Irregular rhythm notification | Good HR detection for sinus rhythm, poor for atrium fibrillation. Irregular rhythm notification acceptable for guiding decisions regarding atrial fibrillation patients | 0 |  |  | 2024-05-01 | 3.7 | 6.7
006_sia_wd_bit | Bittium | Bittium Faros 360 | https://www.bittium.com/medical/bittium-faros | 2019-01-01 | Current | Clinical | not allowed to share price details publicly | CPU + electrodes | Chest | 18 | L48xW13xH29 | 0 | 1; spot electrodes; 3 channels; lead I; lead II; aVF; 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 1000; sternum | 0 | 0 | 0 | 0 | 1 | 168; hours | device charger | 90; minutes | 0 | 0 | 1 | HR pause; bradycardia; tachycardia |  | micro-USB; Blt classic |  | Bittium Faros Manager |  | mircoSD-card | 4320; hours; mb |  | 1 | 0 | 1 | No validation | 0 |  | No validation studies found. Used as reference device in three studies | 0 |  |  | 2024-05-01 | 3.7 | 3.3
007_sia_wd_bit | Bittium | Bittium Faros 180 | https://www.bittium.com/medical/bittium-faros | 2019-01-01 | Current | Clinical | not allowed to share price details publicly | CPU + electrodes | Chest | 18 | L48xW13xH30 | 0 | 1; spot electrodes; 1 channel; aVF; 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 100; sternum | 0 | 0 | 0 | 0 | 1 | 168; hours | device charger | 90; minutes | 0 | 0 | 1 | HR pause; bradycardia; tachycardia |  | micro-USB; Blt classic |  | Bittium Faros Manager |  | mircoSD-card | 4320; hours; mb |  | 1 | 0 | 1 | No validation | 0 |  | No validation studies found. Used as reference device in one study | 0 |  |  | 2024-05-01 | 3.7 | 3.3
008_sia_wd_bit | Bittium | Bittium Faros 180L | https://www.bittium.com/medical/bittium-faros | 2022-05-01 | Current | Clinical | not allowed to share price details publicly | CPU + electrodes | Chest | 18 | L48xW13xH31 | 0 | 1; spot electrodes; 1 channel; aVF; 250 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 100; sternum | 0 | 0 | 0 | 0 | 1 | 336; hours | device charger | 90; minutes | 0 | 0 | 1 | HR pause; bradycardia; tachycardia |  | micro-USB; Blt classic |  | Bittium Faros Manager |  | mircoSD-card | 4320; hours; mb |  | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-05-01 | 4 | 3.3
009_sia_wd_our | Oura | Oura Ring Gen3 | https://ouraring.com/ | 2021-11-01 | Current | Consumer | 314; one-time purchase price including a device but not the required monthly app subscription (6 EUR) | Ring | Finger | 5 | sizes 4-15 | 1; transmission & reflection; green, red and infrared; 125; finger | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 50; finger | 0 | 0 | 1; 0.1; finger | 0 | 1 | 168; hours | device charger | 80; minutes | 0 | 1 | 0 | HR; Daily Activity Score; Daily Readiness Score; Daily Stress Score; Daily Sleep Score; Daily SpO2 averages; Session based HR; Session based HRV; Session based motion count; Sleep based HR; Sleep based HRV; Sleep based readiness; Steps | 5min: HR; 1day: activity, readiness, stress, sleep, SpO2; | BLE; WiFi | Windows NA; iOS 14 or higher; Android 8.0 or higher; macOS NA | Oura App; Oura on the web |  | built-in storage | 168; hours; mb | within the geographic region where data is collected | 1 | 0 | 1 | external | 4 | HR; RMSSD; AVNN; pNN50; SDNN; HF; LF; LF:HF ratio | High validity for HR and RMSSD and acceptable accuracy for nocturnal AVNN, pNN50, HF and SDNN in average-per-night tests with high error rates for LF and LF:HF ratio. | 1 | Participant compliance was high over several weeks. |  | 2024-09-24 | 4 | 7.3
010_sia_wd_min | MindWare Technologies | MindWare Mobile | https://www.mindwaretech.com/mindware-mobile/ | 2014-01-01 | Current | Research | 5499; electrodes and router comes at an extra price this is before tax (21%) | CPU + electrodes | Hip | 370 | L117xW79xH47 | 0 | 1; spot electrodes; 1 channel; 500 | 1; spot electrodes; 500; thorax | 1; 500; facial muscles | 1; spot electrodes; impedence-derived; 500; thorax | 1; spot electrodes; 500; thenar-hyperthenar | 0 | 0 | 1; 3 axes; 50; hip | 0 | 0 | 0 | 1; EOG | 0 | 48; hours | device charger | minutes | 0 | 0 | 1 | ACC; CO; DZ/dt; EMG; LVET; HR; PEP; PTT; Resp; RMSSD; RSA; SCL; Z0 | 1min: all | WiFi; SD-card |  | BioLab |  | SD-card | 48; hours; 4000; mb |  | 1 | 0 | 1 | No validation | 0 |  | No validation studies found. Used as reference device in one study | 0 |  |  | 2024-06-27 | 7 | 2.7
011_sia_wd_nov | Novacor | Diasys 3 plus | https://novacor.com/product/diasys-3-plus/ | 2020-01-01 | Current | Research | 2000; 1760 for the unit and 300 for the associated required software | CPU + strap | Hip | 195 | L85xW85xH25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; auscultatory and oscillometry; upper arm | 1; 3 axes; hip | 0 | 0 | 0 | 0 | 0 | hours | disposable batteries | 0; minutes | 0 | 1 | 0 | BP (mean); DBP; SBP; HR; PP | Manual/scheduled initiation of measurement | Cable USB-C | Windows NA; iOS NA; Android NA; macOS NA | HolterSoft Ultima ABPM | HolterSoft Ultima | built-in storage | hours; mb |  | 1 | 0 | 1 | external | 1 | DBP; SBP | The device is validated according to the AMI/ANSI/ISO 81060-2:2013 standard, ESH-IP revision 2010 and MEDDEV 2.7/1 rev 4 guidelines; and passes all criteria for both the DBP and SBP measures (i.e., for auscullatory and oscillometric measurements). | 0 |  |  | 2024-06-26 | 6.3 | 3
012_sia_wd_car | Carré Technologies Inc. | Hexoskin Proshirt | https://www.hexoskin.com/ | 2019-01-01 | Current | Research & Consumer | 600 | CPU + vest | Chest | 75 | Shirt: 2XS-4XL; CPU: L42xW72xH13 | 0 | 1; touch; 1 channel; lead I; 256 | 0 | 0 | 1; Respiratory Inductance Plethysmography; 2 belts; 128; rib cage and abdomen | 0 | 0 | 0 | 1; 3 axes; 64; near abdomen | 0 | 0 | 0 | 0 | 1 | 36; hours | device charger | 90; minutes | 0 | 1 | 1 | HR;RR;VM;VT;Steps;EE;Cadence;VO2max | 1min: HR;RR;VM;VT;Steps;EE;Cadence;VO2max | Own cable of the system (for transferring all collected data to the computer application); Bluetooth | Windows 7-8; iOS 7 or higher; Android 4.0.4 or higher; macOS 10.8 or higher | Hexoskin OneSync; Hexoskin Dashboard; HxConvertSourceFile | VivoSense is mentioned to work with Hexoskin data, in case the researchers would like a user-friendly interface to process the raw signal data | built-in storage | 600; hours; mb | USA | 1 | 0 | 1 | external | 12 | HR;RR;VT;VM;EE;Steps;HMI | Strong evidence is present that respiratory rate (RR) and heart rate (HR) can both be validly estimated with Hexoskin Proshirt during a variety of light, moderate, and intense activities, whether controlled or naturalistic. Only two studies one of which were naturalistic showed HR to not correspond with reference device values, and a study had a negative conclusion for RR. The reference devices used to compare the Hexoskin parameters were high-quality too. Overall, the accuracy of HR and RR at high intensity exercises decreased. Multiple studies showed test-retest reliability of HR and RR were high across 3-4 days. Most studies suggest tidal volume and minute volume (as well as inspiratory capacity and inspiratory reserve volume) also have acceptable validity, but this significantly decreases with more intense physical activity. Several studies found the tidal volume and minute volume to be out of alignment with a reference device. Energy expenditure (EE) was found to be invalid in all studies that assessed it. Findings for the accuracy of Steps were very mixed. While approximately half of studies that assessed it determined it is valid, other half of studies found it to be inaccurate. Hip Motion Intensity (HMI) was found to be sufficiently valid in one study. | 4 | Hexoskin is perceived by participants to be highly comfortable when used in a laboratory procedure. It can be used in those with chronic obstructive pulmonary disease when the use duration is several hours. Hexoskin can be preferred over wearable watches in studies involving individuals with dementia as the vest cannot be easily removed. Hexoskin use in an ambulatory study of a week, however, is problematic. Combined with an insufficient intrinsic motivation to wear and the need to repeatedly wet the electrocardiogram sensors, older individuals found it too burdensome to be worn in daily life for longer than a day. |  | 2023-11-28 | 7.7 | 3.7
013_sia_wd_app | Apple | Apple Watch SE | https://www.apple.com/apple-watch-se/ | 2020-09-01 | Current | Consumer | 240; iPhone is required | Watch | Wrist | 36 | L44xW38xH10.7 | 1; reflection; green and infrared; wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 1; 3 axes; wrist | 1 | 0 | 1; Microphone; Altimeter; Ambient Light Sensor | 1 | 18; hours | device charger | 90; minutes | 1 | 1 | 0 | HR;HRV-SDNN;Active Time;Steps;EE;Sleep Stages;Sleep Detection;Cadence;Distance;Menstrual Cycle Tracking;Sleep Latency; Sleep Efficiency | 1hour: HR, HRV-SDNN, steps, acitivyt time; 1day: HR, HRV-SDNN, Activity time, Distance, Steps, Sleep Detection, Sleep Stages, Sleep Latency | Blt 5.3; WiFi 4 | Windows NA; iOS iPhone 8 or later with iOS 16 or higher; Android NA; macOS NA | Apple Watch App |  | built-in storage; data has to be transmitted to the coupled Apple Health App (an iPhone is necessary) | hours; 32000; mb | Denmark for Europe, and 5 different server locations for the U.S.: Nevada, North Carolina, Arizona, Oregon, California | 1 | 1 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-07-09 | 2.3 | 4.2
014_sia_wd_app | Apple | Apple Watch Ultra 2 | https://www.apple.com/apple-watch-ultra-2/ | 2023-09-01 | Current | Consumer | 899; iPhone is required | Watch | Wrist | 61 | L49xW44xH14.4 | 1; reflection; green and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 1; 3 axes; wrist | 1 | 1; 0.01; wrist | 1; Barometer; Altimeter; Ambient light sensor; Water temperature;Depth gauge;Microphone | 1 | 36; hours | device charger | 90; minutes | 1 | 1 | 0 | HR;HRV-SDNN; Active Time;Steps;EE;Sleep Stages;Sleep Detection;Cadence;Distance;Menstrual Cycle Tracking;Sleep Latency; Sleep Efficiency; Afib detection; sinus rhythm detection | 1hour: HR, HRV-SDNN, steps, acitivyt time; 1day: HR, HRV-SDNN, Activity time, Distance, Steps, Sleep Detection, Sleep Stages, Sleep Latency | Blt 5.3; WiFi 3 | Windows NA; iOS iPhone Xs or later with iOS 17 or higher; Android NA; macOS NA | Apple Watch App |  | built-in storage; data has to be transmitted to the coupled Apple Health App (an iPhone is necessary) | hours; 64000; mb | Denmark for Europe, and 5 different server locations for the U.S.: Nevada, North Carolina, Arizona, Oregon, California | 1 | 1 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-03-24 | 2.3 | 4
015_sia_wd_app | Apple | Apple Watch Series 9 | https://www.apple.com/apple-watch-series-9/ | 2023-09-01 | Current | Consumer | 449; iPhone is required | Watch | Wrist | 32 | L41xW35xH10.7 | 1; reflection; green and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 1; 3 axes; wrist | 1 | 1; 0.01; wrist | 1; Barometer; Altimeter; Ambient light sensor;Microphone | 1 | 18; hours | device charger | 75; minutes | 1 | 1 | 0 | HR;HRV-SDNN; Active Time;Steps;EE;Sleep Stages;Sleep Detection;Cadence;Distance;Menstrual Cycle Tracking;Sleep Latency; Sleep Efficiency; Afib detection; sinus rhythm detection | 1hour: HR, HRV-SDNN, steps, acitivyt time; 1day: HR, HRV-SDNN, Activity time, Distance, Steps, Sleep Detection, Sleep Stages, Sleep Latency | Blt 5.3; WiFi 4 | Windows NA; iOS iPhone Xs or later with iOS 17 or higher; Android NA; macOS NA | Apple Watch App |  | built-in storage; data has to be transmitted to the coupled Apple Health App (an iPhone is necessary) | hours; 64000; mb | Denmark for Europe, and 5 different server locations for the U.S.: Nevada, North Carolina, Arizona, Oregon, California | 1 | 1 | 1 | No validation | 0 |  | No validity/reliability studies were conducted for this version. See Apple Watch Series 8. | 0 |  |  | 2024-03-24 | 3.3 | 5.3
016_sia_wd_app | Apple | Apple Watch Series 8 | https://support.apple.com/en-us/111848 | 2022-09-01 | Discontinued | Consumer | 400; iPhone is required | Watch | Wrist | 42 | L45xW38xH10.7 | 1; reflection; green and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 1; 3 axes; wrist | 1 | 1; 0.01; wrist | 1; Barometer;Altimeter; Ambient light sensor;Microphone | 1 | 18; hours | device charger | 75; minutes | 1 | 1 | 0 | HR;HRV-SDNN; Active Time;Steps;EE;Sleep Stages;Sleep Detection;Cadence;Distance;Menstrual Cycle Tracking;Sleep Latency; Sleep Efficiency; Afib detection; sinus rhythm detection | 1hour: HR, HRV-SDNN, steps, acitivyt time; 1day: HR, HRV-SDNN, Activity time, Distance, Steps, Sleep Detection, Sleep Stages, Sleep Latency | Bluetooth 5.3 and WiFi | Windows NA; iOS iPhone 8 or later with iOS 16 or later; Android NA; macOS NA | Apple Watch App |  | built-in storage; data has to be transmitted to the coupled Apple Health App (an iPhone is necessary) | hours; 32000; mb | Denmark for Europe, and 5 different server locations for the U.S.: Nevada, North Carolina, Arizona, Oregon, California | 1 | 1 | 1 | external | 4 | SpO2; Sleep time; Awake time; Sleep Stages; HR; Sleep Latency; Sleep Efficiency | Blood oxygen saturation (Sp02) was validated across a gold standard and had good accuracy, with a less than 4% deviation from true values. When two watches were worn at the same time by the same individual (same arm) at upper arm and wrist, heart rate measurements were in agreement, supporting device reliability. However, the test-retest reliability over time has not been investigated. The sleep duration and awake time durations had good validity as compared to polysomnography. The sleep stage detection validity is considered around acceptable (F1 score over 4 sleep stages was .49) though the Apple Watch 8 has been overperformed by all other wearables in a study that compared to polysomnography. Sleep latency was found to be very precise. Sleep efficiency also demonstrated acceptable to good validity. | 2 | Wearing the watch at the upper-arm by surgeons during surgery was shown to be feasible. A few participants who wore the watch and answered daily ecological prompts about the watch use states using less health/physiology oriented aspects of the watch (e.g., checking emails, listening to music) more often. They stated such functions do help with wearing the watch more. They did not see much point in wearing the watch while not working/exercising. However, these usability studies did not have sufficient sample sizes (1-4 people) and should be taken with a grain of salt. |  | 2024-08-27 | 3.7 | 6
017_sia_wd_bio | BIOPAC | Research Ring | https://www.biopac.com/product/research-ring/ | 2023-07-01 | Current | Research | 2020; one-time purchase price of a device, option to interface up to 9 rings with AcqKnowledge Ring software (€5284) | Ring | Finger | 9 | sizes 6-13 | 1; reflection; green; 400; finger | 1; Manually Initiated Recording; 1 channel; lead I; 800 | 0 | 0 | 0 | 1; gold plated dry electrodes; 400; finger | 0 | 0 | 0 | 0 | 0 | 1; 20; finger | 0 | 1 | 10; hours | device charger | 120; minutes | 0 | 1 | 1 | EDRindex, EDRamp, EDL, ΔEDL, HR, HRV, RSA | 1min: EDRindex, EDRamp, EDL, ΔEDL, HR, HRV, RSA | Bluetooth Low-Energy | Windows NA; iOS iOS 15.2 or higher; Android NA; macOS NA | Senstream App | AcqKnowledge Ring software from Biopac | stream and store data via the app - either locally on iPhone or sent to Senstream cloud; via AcqKnowledge - on your PC | 0; hours; 0; mb | USA | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-01-26 | 4 | 3.3
018_sia_wd_now | NOWATCH | NOWATCH | https://nowatch.com/ | 2023-01-01 | Current | Consumer | 450 | Watch | Wrist | 50 | D37xH10.9 | 1; reflection; green, red, infrared; 32; wrist | 0 | 0 | 0 | 0 | 1; dry electrode; wrist | 0 | 0 | 1; 32; wrist | 0 | 0 | 1; 1.67E-2; wrist | 1; Barometer | 1 | 144; hours | Device charger | minutes | 1 | 1 | 1 | HR; IBI; RR; activity count; activity type; cadence; sleep stages; steps | 1min: all | Bluetooth | Windows NA; iOS iOS 15.0 or higher; Android Android 10 or newer; macOS NA | NOWATCH app |  | built-in storage | 24; hours; mb | Europe | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-03-15 | 3.7 | 6
019_sia_wd_goo | Google | Pixel Watch 1 | https://store.google.com/gb/config/google_pixel_watch?hl=en-GB&selections=eyJwcm9kdWN0RmFtaWx5IjoiWjI5dloyeGxYM0JwZUdWc1gzZGhkR05vIn0%3D | 2022-10-01 | Discontinued | Consumer | 325 | Watch | Wrist | 36 | D41xH12.3 | 1; reflection; green, red, and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; wrist | 1; wrist | 1 | 0 | 1; Compass; Altimeter; Ambient Light Sensor | 1 | 24; hours | device charger | 110; minutes | 1 | 1 | 0 | HR; HRV; Breathing rate; Resting HR; Afib notifications; Daily Readiness score; Active Zone Minutes; Cardio Fitness Score; Exercise tracking; Stress Management score; Sleep score; Sleep time; Sleep stage detection | 1min: Blood oxygen saturation during sleep; 1hour: HR, Hourly activity, Stress management score, Resting HR; 1day: Sleep time, Sleep stage detection, RR during sleep, HRV, Skin temp, Steps | Bluetooth 5.0 | Windows NA; iOS NA; Android Android 8.0 or newer; macOS NA | Google Pixel Watch app, Fitbit app | Fitbit ECG app | built-in storage | hours; 32000; mb | several countries within the European Economic Area, United Kindom and Switzerland, controlled by Fitbit International Limited, an Irish company. | 1 | 1 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-03-15 | 2.7 | 4
020_sia_wd_goo | Google | Pixel Watch 2 | https://store.google.com/nl/product/pixel_watch_2?hl=nl | 2023-10-01 | Discontinued | Consumer | 399; Some features require Fitbit Premium (first 6 months free, after that 8.99EUR a month or 79.99EUR yearly) | Watch | Wrist | 31 | D41xH12.3 | 1; reflection; green, red, and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 1; dry electrode; wrist | 0 | 0 | 1; 3 axes; wrist | 1; wrist | 1 | 1; wrist | 1; Compass; Altimeter; Ambient Light Sensor; Barometer; Magnetometer | 1 | 24; hours | device charger | 75; minutes | 1 | 1 | 0 | HR; HRV; Breathing rate; Resting HR; Afib notifications; Daily Readiness score; Active Zone Minutes; Cardio Fitness Score; Exercise tracking; Stress Management score; Sleep score; Sleep time; Skin temp; SpO2; EDA; | 1min: Blood oxygen saturation during sleep; 1hour: Steps; HR (every 2 hours), Hourly activity, Stress management score, Resting HR; 1day: Sleep time, Sleep stage detection, RR during sleep, HRV, Skin temp, Daily blood oxygen saturation, Mean resting HR | Bluetooth 5.1 | Windows NA; iOS NA; Android Android 9.0 or newer; macOS NA | Google Pixel Watch app, Fitbit app | Fitbit ECG app, Fitbit EDA Scan app | built-in storage | hours; 32000; mb | several countries within the European Economic Area, United Kindom and Switzerland, controlled by Fitbit International Limited, an Irish company. | 1 | 1 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-03-15 | 4.3 | 6.3
021_sia_wd_cor | Corsano | Corsano Cardiowatch 287-2 | https://corsano.com/shop/ | 2022-10-01 | Current | Clinical | 1080; the price includes the watch, a blood pressure cuff, and platform access for 1 month - note that portal setup and configuration always comes with a one-off 2000 euro payment- subscription to the portal for a single bracelet over a year is 600 euros | Watch | Wrist | 19 | L25xW40xH10 | 1; reflection; green, red, and infrared; 32; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I; 256 | 0 | 0 | 0 | 1; dry electrode; wrist | 0 | 1; Cuffless Non-Invasive Blood Pressure algorithm; wrist | 1; 3 axes; 32; wrist | 0 | 0 | 1; 0.3; wrist | 0 | 1 | 264; hours | device charger | minutes | 0 | 0 | 1 | Activity; DBP; SBP; Core body temp; EDA; activity ECG; HRV; NIBP; HR; RR; Core body temp; Skin temp; Sleep; Spirometry; SpO2; Steps | 1min: all | Bluetooth BLE 5.0 | Windows NA; iOS NA; Android NA; macOS NA | Corsano Trials App, Medical cloud storage | Corsano Research Portal | built-in storage | 256; hours; mb | Corsano medical cloud storage in Switzerland or Germany | 1 | 1 | 1 | No validation | 0 |  | NA. See the synthesis for the previous version, Corsano Cardiowatch 287-1. | 0 |  |  | 2024-06-26 | 5 | 6.7
022_sia_wd_cor | Corsano | Corsano Cardiowatch 287-1 | https://corsano.com/shop/ | 2021-11-01 | Current | Clinical |  | Watch | Flexible | 15 | L36xW25xH9 | 1; reflection; green, red, and infrared; 32; wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 32; wrist | 0 | 0 | 0 | 0 | 1 | 168; hours | device charger | minutes | 0 | 0 | 1 | Activity; HR; HRV; RR; IBI; Afib; steps; cadence; speed; energy expenditure; calorid data; exercise type; sleep classifications | 1min: all | Bluetooth BLE | Windows NA; iOS NA; Android NA; macOS NA | Corsano Trials App, Medical cloud storage | Corsano Research Portal | built-in storage | 64; hours; mb | Corsano medical cloud storage in Switzerland or Germany | 1 | 1 | 1 | external | 1 | HR; IBI | PPG signal can determine HR and IBI with high accuracy in resting position (no speaking) in cardiovascular at-risk patient (also in the different subgroups for skin type, hair density, age, BMI and gender). | 0 |  |  | 2024-06-26 | 4.3 | 6
023_sia_wd_bio | Biospace Inc. | InBody BAND 2 | https://inbodyusa.com/products/band2/ | 2017-09-01 | Discontinued | Consumer | 149 | Watch | Wrist | 20 | L25.4xW17.78xH10.16 | 0 | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 0 | 0 | 0 | 1; Bioelectrical impedance | 1 | 336; hours | device charger | 60; minutes | 0 | 1 | 0 | %BF; Body Fat Mass; Skeletal Muscle Mass; HR; Steps; EE; Sleep Amount;Sleep Score;Sleep Stages;Active Time;Distance;Activity Detection | 1hour: Steps, EE, HR; 1day: Weight, %BF, Body Fat Mass, Skeletal Muscle Mass, Sleep amount, Duration of each sleep stage | Bluetooth | Windows NA; iOS 15.0 or higher; Android 10.0 or higher; macOS NA | InBody Wear App |  | built-in storage | 336; hours; mb |  | 1 | 0 | 1 | external | 1 | %BF | %BF (percentage body fat) can be measured both with excellent validity and reliability as measured in an adult population (mean age = 45). Other parameters have not been tested yet | 0 |  |  | 2023-12-03 | 2.7 | 3.7
024_sia_wd_fit | Fitbit | Sense | https://www.fitbit.com/global/us/products/smartwatches/sense | 2020-09-01 | Discontinued | Consumer | 150; Some features require a subscription app (first 6 months free, after that 8.99EUR a month or 79.99EUR yearly) | Watch | Wrist | 46 | L40.4xW40.4xH12.4 | 1; reflection; green, red, and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 1; Manually Iinitiated Recording; dry electrode; wrist | 0 | 0 | 1; 3 axes; wrist | 1; 3 axes; wrist | 1 | 1; wrist | 1; Ambient light; altimeter; | 1 | 144; hours | device charger | 120; minutes | 1 | 1 | 0 | HR; HRV; ns.SCR-MIR; Skin temp; Night SpO2; Night RR; Resting HR; Afib notifications; Daily Readiness score; Active Zone Minutes; Cardio Fitness Score; Steps; Distance; Calories; Exercise tracking; Wellness report; Stress Management score; Sleep log; Sleep Stages; Sleep Score | 1min: HR; 1hour: Distance, Calories, Hourly Activity, Exercise Count, Steps, Stress score, Skin temp; 1day: Night RR, HRV during night | Bluetooth 5.0; NFC | Windows NA; iOS iOS 15.0 or higher; Android Android 10.0 or higher; macOS NA | Fitbit App | Fitbit ECG app; Fitbit EDA Scan app | built-in storage | 168; hours; 4000; mb | several countries within the European Economic Area, United Kindom and Switzerland, controlled by Fitbit International Limited, an Irish company. | 1 | 1 | 1 | external | 5 | HR; SCL; Energy Expenditure (EG); Total sleep time (TST); Wake after sleep onset (WASO); sleep efficiency (SE) | Sufficient agreement for HR was reported by two studies under both dynamic and static conditions. However, one study found opposite results, with HR having proper agreement with a reference device for only two out of twelve participants. While skin conductance level was validly estimated, energy expenditure was not. Additionally, the device seems to overestimate Sleep Duration and Sleep Efficiency while underestimating Wake After Sleep Onset. | 0 |  |  | 2024-03-18 | 4.3 | 7.3
025_sia_wd_fit | Fitbit | Sense 2 | https://www.fitbit.com/global/us/products/smartwatches/sense2?sku=521BKGB | 2022-09-01 | Current | Consumer | 223; Some features require a subscription app (first 6 months free, after that 8.99EUR a month or 79.99EUR yearly) | Watch | Wrist | 38 | L40.5xW40.5xH12.3 | 1; reflection; green, red, and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 1; dry electrode; wrist; palm | 0 | 0 | 1; 3 axes; wrist | 1; 3 axes; wrist | 1 | 1; wrist | 1; Ambient light; altimeter; | 1 | 144; hours | solar panel in device; Device charger | 120; minutes | 1 | 1 | 0 | HR; HRV; EDA (both scan-based and continuous); Skin temp; SpO2; Breathing rate; Resting HR; Afib notifications; Daily Readiness score; Active Zone Minutes; Cardio Fitness Score; Steps; Distance; Calories; Exercise tracking; Wellness report; Stress Management score; Sleep Score; Sleep Score Cumulative; | 1min: HR; 1hour: ns.SCR, Distance, Calories, Hourly Activity, Exercise Count, Steps, Stress score, Skin temp; 1day: Night RR, HRV during night | Bluetooth 5.0; NFC | Windows NA; iOS iOS 15.0 or higher; Android Android 10.0 or higher; macOS NA | Fitbit App | Fitbit ECG app; Fitbit EDA Scan app | built-in storage | 168; hours; 4000; mb | several countries within the European Economic Area, United Kindom and Switzerland, controlled by Fitbit International Limited, an Irish company. | 1 | 1 | 1 | No validation | 0 |  | No validation studies yet. See results for previous version (Fitbit Sense). | 0 |  |  | 2024-03-18 | 5.7 | 8
026_sia_wd_gar | Garmin | vivoactive 4 | https://www.garmin.com/en-US/p/643382 | 2019-09-01 | Discontinued | Consumer | 299 | Watch | Wrist | 50 | D40xH12.7 | 1; reflection; green, red, and infrared; wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 1; wrist | 0 | 0 | 1; Blood oxygen saturation sensor; Glonass; Galileo; Compass; Ambient light; Ambient temperature sensor; barometric altimeter | 1 | 192; hours | device charger | minutes | 1 | 1 | 0 | HR; HR zones; SpO2; VO2max; Active Time; Sleep duration; Steps; Stress score; Recovery time; Body Battery; Speed; Pace; Cadence; RR; Sleep stage detection (no rem) | 1min: HR; 5mins: Stress score (around 5 to 30mins per score, difficult to say as app shows data over 15 hours in a small graph) 1hour: Body battery, RR (day and night); 1day: Steps, Intensity minutes, Steps, Calories, Average resting HR, Sleep duration, Sleep stage duration (deep,light, awake, no rem); Notes: if an exercise session is manually started, the activity (including pace and cadence) is recorded per minute | Bluetooth, ANT+, Wi-Fi | Windows NA; iOS iOS 15.0 or higher; Android Android 7 or higher; macOS NA | Garmin Connect | Garmin Connect IQ Store; Garmin Dive; Garmin Golf; Garmin Jr.; Garmin ECG | built-in storage | 200; hours; 4000; mb | several countries, personal data regarding individuals who reside in a country in the European Economic Area, United Kindom or Switzerland is controlled by Garmin Würzburg GmbH; Data from customers in mainland China, is store in GarminChina Shanghai RHQ Co., Ltd's servers in mainland China; Others data will be stored on servers in United States of America, United Kingdom and/or Australia. | 1 | 1 | 1 | external | 4 | SDNN, HRVi, RMSSD, VLF, LF; Classification of inactive, light, moderate-to-vigorous physical activity and raw accelerometry signal (accelerometer); sleep duration | HRV measures - the one VRU study showed excellent accuracy for very low frequency HRV measures, but only moderate to inaccurate for high frequency metrics (especially in patients with cardio- or cerebrovascular diseases), in resting conditions (in the lab). Accelerometer - the validity and reliability of the raw accelerometer data shows neutral to positive results. The estimation of activity levels (inactive, light and moderate to vigorous physical acivity) appears promising, even though the results of raw accelerometer signal compared to a gold standards are moderate. Total sleep duration - the detected TSD showed poor correlation with the sleep diaries. | 0 |  |  | 2024-08-07 | 3.7 | 6.3
027_sia_wd_gar | Garmin | vivoactive 5 | https://www.garmin.com/en-US/p/1057989#specs | 2023-09-01 | Current | Consumer | 300 | Watch | Wrist | 26 | D42.2xH11.1 | 1; reflection; green, red, and infrared; wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 0 | 1 | 0 | 1; Blood oxygen saturation sensor; Glonass/Galileo; Compass; Ambient light | 1 | 264; hours | device charger | 120; minutes | 1 | 1 | 0 | HR; Resting HR; HR zones; HRV; SpO2; VO2max; Active Time; Sleep duration; Stress score; Recovery time; Body Battery; Stress; Speed; Pace; Cadence; RR; Sleep recommendations; RR; Sleep stage detection | 1min: HR; 5mins: Stress score (around 5 to 30mins per score, difficult to say as app shows data over 15 hours in a small graph) 1hour: Body battery, RR (day and night), Distance; 1day: HRV (nighttime), Steps, Intensity minutes, Steps, Calories, Average resting HR, Sleep duration, Sleep stage duration (deep,light, awake, rem); Notes: if an exercise session is manually started, the activity (including speed and distance) is recorded per minute | Bluetooth, ANT+, Wi-Fi | Windows NA; iOS iOS 15.0 or higher; Android Android 7 or higher; macOS NA | Garmin Connect | Garmin Connect IQ Store; Garmin Dive; Garmin Golf; Garmin Jr.; Garmin ECG | built-in storage; data has to be transmitted to the coupled Garmin app. Data can be transmitted to your local computer storage | 264; hours; 4000; mb | several countries, personal data regarding individuals who reside in a country in the European Economic Area, United Kindom or Switzerland is controlled by Garmin Würzburg GmbH; Data from customers in mainland China, is store in GarminChina Shanghai RHQ Co., Ltd's servers in mainland China; Others data will be stored on servers in United States of America, United Kingdom and/or Australia. | 1 | 1 | 1 | external | 1 | GNNS | GNNS scores acceptably well on mountain terrain, in study with one participant. | 0 |  |  | 2026-01-07 | 3.3 | 6.8
028_sia_wd_who | WHOOP | WHOOP 4.0 | https://www.whoop.com/ | 2021-09-01 | Current | Consumer | 30; monthly subscription for a device | Adjustable strap | Wrist | 28 | L43.18XW27.9XH10.16 | 1; reflection; green, red, and infrared; flexible | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; flexible | 1; 3 axes; flexible | 0 | 1; flexible | 1; Pulse oximeter | 1 | 120; hours | Device charger; the device charger itself can be charged using a usb-c connector, and the device charger can then be clicked on the WHOOP itself for charging | 150; minutes | 1 | 1 | 0 | HR; Resting (sleep) HR; Blood oxygen saturation; Daily HRV-RMSSD during sleep; Monthly HRV-RMSSD during sleep; Daily RR during sleep; Averaged Monthly RR during sleep; Number of activities; Percentage daily strain; Activity type classification; Calories; Daily % Recovery; Sleep duration; Active time; Monthly maximal HR; Monthly averaged HR; Monthly Calories; Monthly averaged Sleep duration; Daily Bed Time; Averaged Monthly Bed Time; Daily Total Sleep Disturbances; Averaged Monthly Sleep Disturbances; Nightly Needed Sleep; Sleep Stages; Daily Skin Temperature; Stress score; Sleep performance | 1min: HR; 1hour: Stress score; 1day: Resting (sleep) HR, Blood oxygen saturation, Daily HRV-RMSSD during sleep, Daily RR during sleep, Number of activities, Percentage daily strain, Activity type classification, Calories, Daily % Recovery, Sleep duration, Active time, Daily Bed Time, Daily Total Sleep Disturbances, Nightly Needed Sleep, Sleep Stages, Daily Skin Temperature | Bluetooth 4.2 or higher; WiFi | Windows NA; iOS iOS 14 or higher, Android OS 10.0 or higher; Android iOS 14 or higher, Android OS 10.0 or higher; macOS NA | WHOOP App |  | built-in storage; the data should be synced to the app in mobile phone via Bluetooth | 336; hours; mb | for the European Economic Area, Switzerland, and UK, the data is stored in accordance with GDPR, exact server locations are not provided, the website notes for Taiwan residents, personal data may be stored in and transferred between any of the servers in the United States, Germany, Japan, United Kingdom, France, Canada, and India | 1 | 0 | 1 | No validation | 0 |  | No validation study of this version exists. See studies on the previous version of the device. | 0 |  |  | 2024-08-26 | 4 | 6.7
029_sia_wd_who | WHOOP | WHOOP 3.0 | https://shop.whoop.com/en-eu/collections/whoop3-0 | 2019-01-01 | Discontinued | Consumer | 30; monthly subscription for a device | Adjustable strap | Wrist | 18 | L43.18XW27.9XH10.17 | 1; reflection; green; flexible | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; flexible | 1; 3 axes; flexible | 0 | 0 | 0 | 1 | 120; hours | Device charger; the device charger itself can be charged using a usb-c connector, and the device charger can then be clicked on the WHOOP itself for charging | 150; minutes | 0 | 1 | 0 | HR; Resting (sleep) HR; Blood oxygen saturation; Daily HRV-RMSSD during sleep; Monthly HRV-RMSSD during sleep; Daily RR during sleep; Averaged Monthly RR during sleep; Number of activities; Percentage daily strain; Activity type classification; Calories; Daily % Recovery; Sleep duration; Active time; Monthly maximal HR; Monthly averaged HR; Monthly Calories; Monthly averaged Sleep duration; Daily Bed Time; Averaged Monthly Bed Time; Daily Total Sleep Disturbances; Averaged Monthly Sleep Disturbances; Nightly Needed Sleep; Sleep Stages; Daily Skin Temperature; Stress score; Sleep performance | 1min: HR; 1hour: Stress score; 1day: Resting (sleep) HR, Blood oxygen saturation, Daily HRV-RMSSD during sleep, Daily RR during sleep, Number of activities, Percentage daily strain, Activity type classification, Calories, Daily % Recovery, Sleep duration, Active time, Daily Bed Time, Daily Total Sleep Disturbances, Nightly Needed Sleep, Sleep Stages, Daily Skin Temperature | Bluetooth |  | WHOOP App |  | built-in storage; the data should be synced to the app in mobile phone via Bluetooth | 72; hours; mb | for the European Economic Area, Switzerland, and UK, the data is stored in accordance with GDPR, exact server locations are not provided, the website notes for Taiwan residents, personal data may be stored in and transferred between any of the servers in the United States, Germany, Japan, United Kingdom, France, Canada, and India | 1 | 0 | 1 | external | 4 | HR;HRV-RMSSD;Sleep time; Awake onset; Sleep onset; Sleep stage detection | Heart rate and heart rate variability (RMSSD) was shown to have good validity during sleep and physical activity that does not involve high gross body movement (e.g., dumbell curve). With exercise involving high gross body movement, the watch becomes misaligned with the reference device readings. Sleep detection as compared to polysomnography has high accuracy. Sleep stage detection, however, only moderately aligns with those stages identified by polysomnography. The heart rate and heart rate variability variations day-to-day are not any higher than what is expected for those values based on past literature, and it was argued that the device therefore does not add extra measurement related errors. | 1 | One usability study asking three individuals using WHOOP to reflect on their device use showed that individuals who are doing the sports that the device is marketed for (i.e. triathletes) are more likely to use the device. The monthly reports sent by the WHOOP can help users reflect on the impact of interventions/exercises they do on their body and recovery. |  | 2024-08-26 | 4 | 6.8
030_sia_wd_cir | Circular | Circular Slim | https://www.circular.xyz/ | 2023-12-01 | Discontinued | Consumer | 263 | Ring | Finger | 2 | sizes 4-16 | 1; reflection; green, red, and infrared; finger | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; finger | 0 | 0 | 1; finger | 1; Pulse oximeter | 1 | 120; hours | device charger | 45; minutes | 1 | 1 | 0 | HR; Morning Resting HR; RR during Sleep; Skin Temp; Stress Score; Energy Level; HRV-RMSSD during Sleep; Intermittent HRV-RMSSD; Steps; Sleep Duration; Sleep Disturbances; Sleep efficiency; Tranquility; Circadian Rhythm; REM sleep; Deep sleep; Time to fall asleep; Sleep debt; Awake times during sleep; Distance; EE;%maxHR; SpO2; Duration of sleep stages | 1min: HR, HRV, %maxHR, Activity intensity, Sp02; RR; 1hour: Energy level, Stress score; 1day: Sleep Score, Duration of sleep stages, Steps, Calories, Active minutes, EE | Bluetooth Smart | Windows NA; iOS iOS 11.0 or later; Android Android 6.0 or later; macOS NA | Circular |  | built-in storage; data has to be synced with the Circular App in smartphone at least every 10 days as the device cannot store for longer than 10 days | 240; hours; mb |  |  | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-03-29 | 2 | 4.5
031_sia_wd_mov | movisens | EcgMove4 | https://www.movisens.com/en/products/ecg-sensor/ | 2021-01-01 | Current | Research | 2089; Includes the CPU and one chest belt | CPU + strap | Chest | 26 | L11.5XW62.3XH38.6 | 0 | 1; 1 channel, Lead I; 1024 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 64; Chest | 1; 3 axes; 64; Chest | 0 | 1; 1; Chest | 1; Barometer | 1 | 72; hours | device charger | 120; minutes | 1 | 1 | 1 | HR; Step Count; RMSSD;EE;HRV;Movement Acceleration;Activity Class;Body Position; Skin Temp; Atmospheric Air Pressure; Angular Rate | 1min: all | USB cable; Bluetooth Smart 4.0 | Windows Windows 7, 8, 8.1, 10; iOS NA; Android NA; macOS NA | SensorManager | DataAnalyzer | built-in storage; the data is transferred from the local device storage to a PC using a USB cable | 336; hours; 4000; mn |  | 1 | 0 | 1 | external | 1 | HR; RMSSD; HF | All heart rate, heart rate variability in the form of RMSSD and in the form of HF (high frequency power) showed great alignment with the gold standard in both a laboratory procedure and in daily life, in an adult sample. Half-split reliability of the HR and HRV indices were high. Parameters derived from the accelerometer are not yet validated. The rate of valid data in ambulatory settings for HRV was higher than 90% and for daily life recordings higher than 80% (minimal amount of low quality data). | 0 |  |  | 2024-06-12 | 5.8 | 8
032_sia_wd_wit | Withings | Scanwatch | https://www.withings.com/eu/en/scanwatch-2 | 2021-11-01 | Current | Consumer | 299.95 | Watch | Wrist | 58 | D38xH13.2 | 1; Reflection; Green light; Wrist | 1; Manually initiated recording; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; Wrist | 0 | 0 | 0 | 0 | 1 | 720; hours | device charger | 120; minutes | 1 | 1 | 0 | Afib-MIR; HR; QRS; HR; QT; QTc; SpO2; activity tracking; sleep quality score; sleep stage light; sleep stage deep; sleep duration; sleep interruptions; sleep regularity; menstrual cycle tracking; Breathing disturbances; steps | 5min: HR; 1day: Sleep Quality Score, Sleep Stage Duration, Sleep SpO2; Daily SpO2, Breathing Disturbances, Activity tracking, Steps | Bluetooth BLE | Windows NA; iOS iOS 14 and higher; Android Android 10.0 and higher; macOS NA | Health Mate app |  | built-in storage | hours; mb | France | 1 | 1 | 1 | external | 10 | Afib; HR; QRS; PR; QT; QTc; SpO2; | HR can generally be accurately detected, with the best result in stationary conditions. AFib classification varies between studies, but has potential to be an accurate detection algorithm. Other ECG parameters such as the QRS, PR, and QTc show generally accurte results. However, one study used the device in different positions to obtain different ECG Leads which increase the amplitude, duration and morphology detections. The four studies evaluating the SpO2 show varying results; some very promising in multiple hypoxia condition, but the largest studies (n = 36-49) show less positive results and urge to interpret SpO2 results with care. | 0 |  |  | 2024-07-05 | 3 | 6.8
033_sia_wd_wit | Withings | Scanwatch 2 | https://www.withings.com/eu/en/scanwatch-2 | 2023-10-01 | Current | Consumer | 349.95 | Watch | Wrist | 35 | D38.4xH13.4 | 1; Reflection; Green, red and infrared light; Wrist | 1; Manually initiated recording; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; Wrist | 0 | 0 | 1; Wrist | 1; Altimeter | 1 | 720; hours | device charger | 120; minutes | 1 | 1 | 0 | Afib-MIR; HR; Nighttime HRV; QRS; HR; QT; QTc; SpO2; activity tracking; sleep quality score; sleep stage light; sleep stage deep; sleep duration; sleep interruptions; sleep regularity; Skin temp; menstrual cycle tracking; Breathing disturbances; steps; RR | 5min: HR; 1day: Sleep Quality Score, Sleep Stage Duration, Sleep SpO2; Daily SpO2, Breathing Disturbances, Activity tracking, Steps, Nighttime HRV | Bluetooth BLE | Windows NA; iOS iOS 15 and higher; Android Android OS 9 and higher; macOS NA | Health Mate app |  | built-in storage | 168; hours; mb | France | 1 | 1 | 1 | No validation | 0 |  | NA (see previous version) | 0 |  |  | 2024-07-05 | 2.7 | 5.7
034_sia_wd_spa | Spacelabs | OnTrak | https://spacelabshealthcare.com/products/diagnostic-cardiology/abp-monitoring/ontrak/ | 2019-01-01 | Current | Research & Clinical | 3249 | CPU + strap | Hip | 172 | L10.15xH2.76xW7.0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; Oscillometry; Upper arm | 0 | 0 | 0 | 0 | 0 | 1 | 168; hours | Disposable batteries | 0; minutes | 1 | 1 | 0 | DBP;SBP;MAP;PP;HR | Manual/scheduled initiation of measurement | Cable | Windows Windows 7 SP1 Professional, Ultimate and Enterprise (64 bit), Windows 10 Professional and Enterprise (64 bit), Windows Server 2012 R2, Windows Server 2016 & 2019; iOS NA; Android NA; macOS NA | Sentinel |  |  | hours; mb |  | 1 | 1 | 1 | external | 1 | SBP; DBP | The Spacelabs 90227 OnTrak device fulfils the requirements of three recognized validation protocols (European Hypertension Society International Protocol revision 2010 (ESH-IP2), the British Hypertension Society Protocol (BHS) , and the American National Standards Institute/Association for the Advancement of Medical Instrumentation/International Standards Organization (ANSI/AAMI/ISO) 81060-2:2009 protocol) in adults and children and can therefore be recommended for clinical use in standard function. | 0 |  |  | 2024-07-12 | 6.7 | 2.3
035_sia_wd_gar | Garmin | Vivosmart 5 | https://www.garmin.com/en-US/p/782585/ | 2022-04-01 | Current | Consumer | 150 | Watch | Wrist | 25 | L19.5xW10.7xH217 | 1; Greend, red and infrared light; Wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; Wrist | 0 | 0 | 0 | 1; Ambient light sensor; blood oxygen saturation sensor | 1 | 168 | device charger | 90 | 1 | 1 | 0 | RR; Sleep score; Blood oxygen during the day; Blood oxygen during sleep; HR;Daily resting pulse rate; Fitness age; Body battery level; Daily stress score;Steps;EE;Distance;Intensity minutes;HR zones; HR alerts; % HR max; HR calories; Duration of each sleeping stage; Incident Detection; Sleep stage classification (light, deep, REM, awake) | 1min: HR, Stress Score, Body Battery; 1hour: RR; 5hours: Blood oxygen saturation; 1day: Sleep score, Daily stress score, Duration of each sleeping stage, Fitness age | Bluetooth Smart; ANT+ | Windows NA; iOS Apple iOS 16.0 and above; Android Android 7.0 OS and above; macOS NA | Garmin Connect App | Garmin Express | built-in storage | 339; hours; mb | data of individuals who reside in the European Economic Area, United Kindom or Switzerland is stored by Garmin Würzburg GmbH; Data from customers in China, is store in GarminChina Shanghai RHQ Co., Ltd's servers in mainland China; Others data is stored on servers in United States of America, United Kingdom and/or Australia. | 1 | 0 | 1 | No validation | 0 |  | The specific version of this device has no validity/reliability studies. See Garmin Vivosmart 4. | 0 |  |  | 2024-08-26 | 3.3 | 7
036_sia_wd_gar | Garmin | Vivosmart 4 | https://www.garmin.com/en-US/p/605739/ | 2018-09-01 | Discontinued | Consumer | 150 | Watch | Wrist | 17 | L15xW10.5xH197 | 1; Greend, red and infrared light; Wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; Wrist | 0 | 0 | 0 | 1; Ambient light sensor; blood oxygen saturation sensor; barometric altimeter | 1 | 168; hours | device charger | 129; minutes | 1 | 1 | 0 | Sleep score; Blood oxygen during the day; Blood oxygen during sleep; HR;Daily resting pulse rate; Fitness age; Body battery level; Daily stress score;Steps;EE;Distance;Intensity minutes;HR zones; HR alerts; % HR max; HR calories; Duration of each sleeping stage; Incident Detection; VO2max; Sleep stage classification (light, deep, REM, awake) | 1min: HR, Body battery, Stress Score; 5min: SpO2%; 1day: Sleep score, Daily stress score, Duration of each sleeping stage, Fitness age | Bluetooth | Windows NA; iOS Apple iOS 16.0 and above; Android Android 7.0 OS and above; macOS NA | Garmin Connect App | Garmin Express | built-in storage | 339; hours; mb | data of individuals who reside in the European Economic Area, United Kindom or Switzerland is stored by Garmin Würzburg GmbH; Data from customers in China, is store in GarminChina Shanghai RHQ Co., Ltd's servers in mainland China; Others data is stored on servers in United States of America, United Kingdom and/or Australia. | 1 | 0 | 1 | external | 4 | HR; sleep estimates (Total sleep time (TST), time in bed (TIB), wake, wake after sleep onset (WASO), sleep onset latency (SOL), sleep stages, REM latency, sleep onset and sleep end); step count; SpO2 | Garmin Vivosmart 4 has shown acceptable agreement with a chest monitor for HR among individuals with severe disabilities participating in the indirect movement program. It was also shown to be sufficiently acurate for estimating sleep duration but not for sleep architecture, for example detecting sleep stages. Intra-device reliability for sleep estimates was assessed in one participant with two devices over 23 nights and was reported as high. While the watch is valid at counting steps at regular walking speeds, it has compromised validity at slow walking speeds among healthy adults and no validity for step counting among Parkinson's patients. The device was also tested with Chronic Obstructive Pulmonary Disease patients on SpO2 estimation showing insufficient validity due to overstimation of SpO2 at low- and underestimation at high saturation levels. | 1 | Garmin Vivosmart (previous version) was one of the devices evaluated in a user acceptence study with non-metrapolitan cancer survivors who wore it for two weeks. The device was seen positively due to its aesthetics, comfort of wear and accessible design with bigger display and font being well perceived by older users. |  | 2025-02-17 | 3.7 | 7.5
037_sia_wd_gar | Garmin | Index BPM Smart Blood Pressure | https://www.garmin.com/en-US/p/716808/ | 2022-09-01 | Current | Consumer | 176 | CPU + strap | Upper arm | 320 | L58xW145xH47 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; Manometer; Upper arm | 0 | 0 | 0 | 0 | 0 | 1 | 6480; hours | Disposable batteries | 0; minutes | 1 | 1 | 0 | DPB;SBP;HR | Manual/scheduled initiation of measurement | Wi-Fi; Bluetooth | Windows NA; iOS NA; Android NA; macOS NA |  | Garmin Connect App | 100 measurements | 100; measurements | data of individuals who reside in the European Economic Area, United Kindom or Switzerland is stored by Garmin Würzburg GmbH; Data from customers in China, is store in GarminChina Shanghai RHQ Co., Ltd's servers in mainland China; Others data is stored on servers in United States of America, United Kingdom and/or Australia. | 1 | 1 | 1 | internal | 1 | SBP;DBP;HR | The internal clinical validation study is not published in a peer-reviewed journal but it helped receive an FDA clearance. In this study, the device can measure both blood pressure (+/- 2%) and pulse rate (+/- 5%) validly as revealed by measurements within the error margins determined by the American Heart Association. As study sample characteristics and task details were not given, neutral verdicts were given for blood pressure and pulse rate monitoring. | 0 |  |  | 2024-08-26 | 4.2 | 2
038_sia_wd_sky | SkyLabs | CART BP | https://www.skylabs.io/en | 2023-03-01 | Current | Clinical & Consumer | 447 | Ring | Finger | 4 |  | 1; Reflection; Green, red and infrared light; 175; Finger | 0 | 0 | 0 | 0 | 0 | 0 | 1; Machine learning model using PPG data; Finger | 0 | 0 | 0 | 0 | 0 | 1 | 48; hours | device charger | 120; minutes | 0 | 0 | 0 | HR; HRV; DBP; SBP; BPV | 5min: HR, DBP, SBP, BPV | Bluetooth | Windows NA; iOS iOS 10.0 and later; Android Android 6.0 and later; macOS NA | CART BP pro |  | built-in storage | 24; hours; mb |  | 0 | 0 | 0 | No validation | 0 |  |  | 0 |  |  | 2024-09-04 | 1.7 | 4
039_sia_wd_ama | Amazfit | Helio Ring | https://us.amazfit.com/products/amazfit-helio-ring | 2024-05-01 | Current | Consumer | 300 | Ring | Finger | 4 | sizes 8, 10, 12 | 1; Reflection; dual-color LED (specific lights types not provided); Finger | 0 | 0 | 0 | 0 | 1; Dry electrode; Finger | 0 | 0 | 1; 3 axes; Finger | 1; 3 axes; Finger | 0 | 1; Finger | 1; Blood oxygen saturation sensor | 1 | 48; hours | device charger | 120; minutes | 1 | 1 | 0 | Readiness score; HR; Blood oxygen saturation; Stress level; Sleep stages; Sleep breathing quality; Daytime naps; Sleep score; HR recovery; EDA; HRV; Emotional response detection; Active minutes; Steps; EE; Physical recovery; Mental recovery; Breathing quality; VO2max; Training load; Nocturnal resting HR | 1min: Speed, Cadence;5min: HR, Altitude;1day: Sleep score, Nocturnal resting HR, Breathing quality, Skin temp, Sleep stage durations (light, deep, REM), Count of awake times during sleep, Daily HR, VO2max, Training load | BLE | Windows NA; iOS iOS 14.0 and above ; Android Android 7.0 and above; macOS NA | Zepp App |  | built-in storage | hours; mb | the server country is not specified, but the company is based in China | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-09-04 | 3 | 5.2
040_sia_wd_bio | BioStrap | Biostrap Kairos | https://shop.biostrap.com/products/biostrap-kairos-pilot-set | 2023-07-01 | Current | Clinical & Consumer | 275 | Adjustable strap | Flexible | 10 | L36xW22xH12.5 | 1; Reflection; two green and 1 infrared light; Wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; Wrist | 1; 3 axes; Wrist | 0 | 0 | 1; Blood oxygen saturation sensor | 1 | 120; hours | device charger | minutes | 0 | 1 | 0 | HR; HRV; RR; SpO2; sleep analysis; steps; activity metrics; arterial and peripheral elasticity | 5min: HR, HRV, SpO2; sleep analysis; steps; activity metrics | Bluetooth 5.0 | Windows NA; iOS NA; Android NA; macOS NA | Vital Science application | Remote physiological monitoring web dashboard | built-in storage | hours; 64; mb |  | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-09-04 | 2.3 | 4.3
041_sia_wd_car | CardioMood | CardioMood 287-2 | https://cardiomood.com/product/cardiowatch-287-2 | 2023-01-01 | Current | Research & Consumer | 499 | Wrist strap | Wrist |  | 40Lx25Hx10W | 1; Green, red and infrared light; 64-128; Wrist | 0 | 0 | 0 | 0 | 1; Dry electrodes; Wrist | 0 | 0 | 1; 3 axes; 64-128; Wrist | 0 | 0 | 0 | 1; Heat flux; Blood oxygen saturation sensor | 1 | 240; hours | device charger | minutes | 0 | 1 | 1 | Steps; Distance; EE; Core Body temp; HRV-MRR; HRV-SDNN; HRV-RMSSD; HRV-PNN50;HRV-MO; HRV-AMO; HRV-SI; IBI; HR; SpO2; Sleep stage classification (light, deep, awake, REM); RR; Performance; Sleep Tranquility score; Sleep Consistency Score; Sleeptime; Resting HR; Readiness score; EDA; Stress | 1mins: HR and HRV indices IF Preventicus heartbeats software is used; 5mins: Readiness score; 10mins: HR; 1hour: EE (calories), steps; 1day: Core body temperature, HRV, RR | Bluetooth 5.0 | Windows NA; iOS iOS 12.1 or higher; Android Android 8.0 or higher; macOS NA | CardioMood Trends | Preventicus heartbeats | built-in storage | hours; 64; mb | Preventicus heartbeats | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2025-02-12 | 2.3 | 4.8
042_sia_wd_sam | Samsung | Galaxy Ring | https://www.samsung.com/us/rings/galaxy-ring/ | 2024-08-01 | Current | Consumer | 449 | Ring | Finger | 3 | sizes 5-15 | 1; Reflection; Green light; Finger | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; Finger | 0 | 0 | 1; Finger | 1; Blood oxygen saturation sensor | 1 | 168; hours | device charger | 80; minutes | 0 | 1 | 0 | HR;SpO2;Sleep duration;Energy score; Workout tracking; Skin Temp; Active sleep score; Snoring analysis; Movement during sleep; Sleep latency; RR; Sleep quality score; Sleep routine score; Min HR; Max HR; Period onset estimation using skin temp | 1hour: SpO2 during sleep, Skin temp, Min HR, Max HR, Movement during sleep;1day: Sleep duration, Sleep stage durations (light, deep, REM, awake), Daily HR, Restfulness score, Mental recovery score, Sleep Score, Physical recovery | BLE 5.4 | Windows NA; iOS NA; Android Android 11 or above; macOS NA | Samsung Health App |  | built-in storage | hours; 8000; mb | servers are located in South Korea, U.S.A., several locations in Europe, among other locations | 1 | 0 | 0 | No validation | 0 |  |  | 0 |  |  | 2024-09-04
043_sia_wd_goo | Google | Pixel Watch 3 | https://store.google.com/product/pixel_watch_3?hl=nl | 2024-01-01 | Current | Consumer | 399 | Watch | Wrist | 31 | D41xH12.3 | 1; Reflection; Green, red and infrared light; Wrist | 1; Manually initiated recording; one channel similar to Lead I | 0 | 0 | 0 | 1; Dry electrodes; Wrist | 0 | 0 | 1; 3 axes; Wrist | 1; Wrist | 1 | 1; Wrist | 1; SpO2; Compass; Altimeter; Ambient Light Sensor; Barometer; Magnetometer | 1 | 36; hours | device charger | 60; minutes | 1 | 1 | 0 | Steps; Distance; Floors; EE; Readiness score; HR; Sleep quality; Cardio load; Target cardio load; HRV; Resting HR; Sleep stages (light, deep, awake REM); Sp02%; Skin temp variation; RR; Stress management score; Physical calmness; Activity balance; | 1min: HR; 1hour: Steps, Distance, EE; 1day: Resting HR, RR, SpO2%, Skin temp variation | Bluetooth 5.0; Ultra-wideband; Wi-Fi 802.11 | Windows NA; iOS NA; Android Android 10.0 or higher; macOS NA | Google Pixel Watch; Fitbi App; Fitbit ECG App |  | built-in storage | hours; 32000; mb |  | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2025-02-13
044_sia_wd_adi | AD Instruments | Equivital eq02+ Lifemonitor | https://equivital.com/products/eq02-lifemonitor | 2021-01-01 | Current | Research | 1977 | CPU + vest | Chest | 38 | 78x55x11 | 0 | 1; Two channels similar to lead I and lead II, based on touch connectivity (vest to skin); 256; Chest | 0 | 0 | 1; Using Strain Gauge Transducer; 25.6; Chest | 0 | 0 | 1; Optional using cuff-based compatibble ancillary sensor (Note, not standard!); Upper arm | 1; 3 axes; 25.6-256; Chest | 0 | 0 | 1; Accuracy +/- 0,1 degree Celsius; measures every 15 seconds; 0.067 | 1; Core temperature; Oximeter | 1 | 48; hours | Device charger; either a dedicated USB compatible cable, or a docking station for charging up to 6 modules | 60; minutes | 1 | 1 | 1 | HR; HR confidence index; IBI; RR; RR confidence index; Skin Temp; Physiological Welfare Index state; Fall state | 15sec: all, 5sec:HR (if configured) | cable; system specific-USB cable; Bleutooth Compliance version 2.1 | Windows NA; iOS NA; Android NA; macOS NA | Equivital™ Manager | Equivital™ Qiosk; eqView Professional (Non-Diagnostic); Equivital™ Black Ghost eqView Mobile; | built-in storage | 1200; hours; 8000; mb |  | 1 | 1 | 1 | external | 4 | HR; RMSSD; SDNN; pNN50; average of NN intervals; Estimated Core Body Temp; Simulated breath hold, simulated chest motion signal linearity | Valid and reliable for estimating core body temperature with a slight overestimation of 0.1°C. It provides reliable HR, RMSSD and pNN50 measurements but only moderate reliability for SDNN and other HRV parameters , with artifacts affecting accuracy. It's respiratory sensor is not accurate for estimating chest movement or breath hold as shown by a simulated benchtop test. | 0 |  |  | 2022-02-05 | 6.8 | 3
045_sia_wd_med | Medtronic Zephyr | BioHarness 3.0 | https://www.zephyranywhere.com/system/overview | 1899-12-31 | Current | Research & Consumer | 766; Including biomodule, one charger and a strap; optionally add GPS for 190 and GPS charger 5 bay for 340, gateway for live monitoring 450; Omnisense software 4530, and SDK kit for downloading and streaming and matlab script 170; all excluding VAT | CPU + electrodes | Chest | 18 | 28Dx7H | 0 | 1; One channel similar to lead I; 256 | 0 | 0 | 1; Measured via belt at rib cage; 18; Thorax | 0 | 0 | 0 | 1; 3 axes; 100; Thorax | 0 | 1; Optional addition can be purchased | 0 | 0 | 1 | 35; hours | device charger | 180; minutes | 0 | 0 | 0 | ACC; RR; HR; HRV; % HR max; % HR at anaerobic threshold; SpO2; posture detection; estimated core temp; impact; activity; caloric burn; activity metrics | 1min: all | cable; system specific-USB cable; Bleutooth Compliance version 2.1 + EDR and version 4.0 | Windows Windows 8 or higher; iOS NA; Android NA; macOS NA | SDK Kit | Omnisense software | built-in storage | 450; hours; mb |  | 1 | 1 | 1 | external | 6 | HR; HRV; RR; accelerometry; step counts | The Zephyr BioHarness 3.0 can reliably measure HR and RR with high test-retest reliability results (72 hours apart). The convergent validity of HR and HRV is examined in merely one individual and shows positive results for a resting condition for HR, whereas HRV showed poor results. The convergent validity of RR is also assessed in merely one individual but showed positive results. The accelerometer signal for step counts shows valid accuracy when compared to video recording; with the most accurate results for jogging, a slight underestimation for walking and small overestimation for running. | 0 |  |  | 2024-08-14 | 6.5 | 5.4