## df_osf_sia_wd
device_id | Manufacturer | Device | Website | Release date | Market status | Main use | Device costs | Wearable type | Location | Weight | Size | PPG | ECG | ICG | EMG | Respiration | EDA | EEG | BP | Accelerometer | Gyroscope | GPS | Skin temperature | Other signals | Water resistance | Battery life | Charging method | Charging duration | Bio-cueing | Bio-feedback | Raw data available | Provided parameters | Parameter sampling window (1min; 5min; 1hour; 1day) | Data transfer method | Compatibility | Required software | Additional software | Internal storage method | Device storage capacity | Server Data Storage | GDPR compliance | FDA approval/clearance | CE approval/label | Highest level of Validation Evidence | Number of validity and reliability studies reviewed | Studied parameters | General validity and reliability synthesis | Number of usability studies reviewed | General usability synthesis | Hyperlink to the device VRU page | Most recent date of RVU search | SiA Expert score (short-term) | SiA Expert score (long-term)
001_sia_wd_emp | Empatica | E4 | https://www.empatica.com/research/e4/ | 2015-01-01 | Discontinued | Research | 1542; one-time purchase price including a device + software | Watch | Wrist | 40 | L44xW40xH16 | 1; reflection; 64; wrist | 0 | 0 | 0 | 0 | 1; dry electrode; 4; wrist | 0 | 0 | 1; 3 axes; 32; wrist | 0 | 0 | 1; 0.2(36-39); 4; wrist | 0 | 1 | 48; hours | device charger | 120; minutes | 0 | 1 | 1 |  |  | BLE; USB 2.0 | Windows 7-10; iOS 8.2 or higher; Android 8.1 or higher; macOS NA | E4 manager; Empatica Connect web dashboard; | E4 Realtime | built-in storage | 60; hours; 60; mb | USA | 1 | 0 | 1 | external | 10 | HF; HR; IBI; ns.SCRs; NS-SCRs-AMP; NS-SCRs-FREQ; RMSSD; S; SCR-AMP; SCRs; SCL; SD1; SD1UP; SD1DOWN; SD2; SDNN ; SDRR | Most studies comparing E4 to a reference device (RD) suggest high validity of HR under static and dynamic conditions (one study found opposite results) and often report HRV parameters as valid but only in static conditions, although results seem conclusive only for RMSSD and SDNN. Movement and speech have been shown to result not only in decreased accuracy but also in high amounts of missing data. Missing data was also observed in static conditions and was attributed to issues with E4’s automatic IBI detection. It also appears E4 cannot provide convergently valid EDA (possibly due to differences in device location). However, one study reported high correlation of SCL with an RD despite it measuring from fingers, with yet another study reporting validity of amplitude of SCRs for strong sustained stressors. | 5 | E4 was generally considered easy to use with charging and handling the data reported as burdensome in multiple studies. Adherence was found to be high but results are inconsistent when it comes to wearing comfort. |  | 2023-06-27 | 5 | 6.7
002_sia_wd_emp | Empatica | EmbracePlus | https://www.empatica.com/embraceplus/ | 2020-01-01 | Current | Research & Clinical | 1984; three year bundle per a device + software (25% academic discount available) | Watch | Wrist | 38 | D32xH15 | 1; transmission & reflection; green, red and infrared; 26-208; wrist | 0 | 0 | 0 | 0 | 1; dry electrode; 1-4; wrist | 0 | 0 | 1; 3 axes; 26-208; wrist | 1; 3 axes; 26-208; wrist | 0 | 1; 0.1(30-45); 1-4; wrist | 0 | 1 | 48; hours | device charger | 90; minutes | 0 | 1 | 1 | HR; HRV; RR; Skin Temp; SpO2; EDA; Activity Counts; Sleep Detection | 1min: HR, HRV; 1day:Activity Counts | BLE; Blt 5.0 | Windows NA; iOS 11 or higher; Android 8.0 or higher; macOS NA | Care app; Care web portal |  | built-in storage | 24; hours; 128; mb |  | 1 | 1 | 1 | internal | 1 | SpO2 | Reliable and valid for use in clinical SpO2 measurement under static conditions, for both male and female subjects as well as for different skin tones. | 0 |  |  | 2023-09-18 | 5.7 | 8
003_sia_wd_vri | Vrije Universiteit van Amsterdam | VU-AMS 5fs | http://www.vu-ams.nl/ | 2007-01-01 | Current | Research | 4550 | CPU + electrodes | Hip | 260 | L126xW69xH32 | 0 | 1; spot electrodes; 2 channels; lead I; lead II; 1000 | 1; spot electrodes; 1000; thorax | 0 | 1; spot electrodes; impedence-derived; 250; thorax | 1; spot electrodes; 10; wrist-thenar | 0 | 0 | 1; 3 axes; 1000; hip | 0 | 0 | 0 | 0 | 0 | 24; hours | disposable batteries | 0; minutes | 0 | 0 | 1 | HR; HRV; RR; Skin Temp; SpO2; EDA; Activity Counts; Sleep Detection | 1min: CO, ECG-Q, HR, IBI, ICG-B, ICG-C, ICG-X, LVET, PEP, RMSSD, RR, RSA, SDNN, SV, TWA | SD-card; Blt | Windows NA; iOS NA; Android NA; macOS NA | VU-DAMS |  | SD-card | hours; 1024; mb |  | 1 | 0 | 1 | external | 3 | IBI; RBI; PEP; LVET; SV; CO; HR | Reliable and valid for measuring cardiac parameters using ICG and ECG. Absolute SV and CO values can deviate, especially during activity. Additionally, the VU-AMS 5fs device has been used as a gold standard reference device for various studies. | 0 | The validation studies showed it is feasible to measure with the VU-AMS device in daily life for 24h. |  | 2023-11-20 | 8.7 | 3
004_sia_wd_vri | Vrije Universiteit van Amsterdam | VU-AMS Core | http://www.vu-ams.nl/ | 2024-03-01 | Current | Research | 5995; one-time purchase price including a device + software | CPU + electrodes | Chest | 52 | L57xW57xH17 | 0 | 1; spot electrodes; 3 channels; lead I; lead II; aVF; 1000 | 1; spot electrodes; 1000; thorax | 0 | 1; spot electrodes; impedence-derived; 250; thorax | 0 | 0 | 0 | 1; 3 axes; 1000; upper sternum | 1; 3 axes; 1000; upper sternum | 0 | 0 | 0 | 0 | 28; hours | device charger | 180; minutes | 0 | 0 | 1 | CO; ECG-Q; HR; IBI; ICG-B; ICG-C; ICG-X; LVET; PEP; RMSSD; RR; RSA; SDNN ; SV; TWA | 1min: all parameters | mircoSD-card | Windows NA; iOS NA; Android NA; macOS NA | VU-DAMS |  | mircoSD-card | 160; hours; 32000; mb |  | 1 | 0 | 0 | No validation | 0 |  | No validation studies yet. This device is the successor of the VU-AMS 5fs device. | 0 |  |  | 2023-11-20 | 7.7 | 4.3
005_sia_wd_fit | Fitbit | Charge 5 | https://www.fitbit.com/global/nl/products/trackers/charge5?sku=421BKBK | 2021-09-01 | Current | Consumer | 150 | Watch | Wrist | 28 | L36.7xW22.7xH11.2 | 1; reflection; wrist | 1; Manually Initiated Recording; 1 channel; lead I | 0 | 0 | 0 | 1; dry electrode; only the scan function, no continuous EDA; wrist | 0 | 0 | 1; 3 axes; wrist | 0 | 0 | 1; wrist | 1; ambient light | 1 | 168; hours | device charger | 120; minutes | 0 | 1 | 0 | EDA-MIR; HR; HRV; RR; Skin Temp; SpO2; Cardio-Fitness Score; Hourly Activity; Sleep Duration; Sleep Score; Sleep Stages; Stress Management Score | 1hour: HR, Steps, Hourly Activity; 1day: Steps, Distance, Floors, Calories burned, Sleep Score; Stress Management Score; Sleep stages | Blt 5.0; NFC | Windows NA; iOS 15 or higher; Android 9.0 or higher; macOS NA | Fitbit app | Fitbit ECG app; Fitbit EDA Scan app | built-in storage; need to connect and send data to the Fitbit app to see the data | hours; mb | several countries within the European Economic Area, United Kindom and Switzerland, controlled by Fitbit International Limited, an Irisch company. | 1 | 1 | 1 | external | 1 | HR, Irregular rhythm notification | Good HR detection for sinus rhythm, poor for atrium fibrillation. Irregular rhythm notification acceptable for guiding decisions regarding atrial fibrillation patients | 0 |  |  | 2024-05-01 | 3.7 | 6.7
006_sia_wd_bit | Bittium | Bittium Faros 360 | https://www.bittium.com/medical/bittium-faros | 2019-01-01 | Current | Clinical | not allowed to share price details publicly | CPU + electrodes | Chest | 18 | L48xW13xH29 | 0 | 1; spot electrodes; 3 channels; lead I; lead II; aVF; 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 1000; sternum | 0 | 0 | 0 | 0 | 1 | 168; hours | device charger | 90; minutes | 0 | 0 | 1 | HR pause; bradycardia; tachycardia |  | micro-USB; Blt classic |  | Bittium Faros Manager |  | mircoSD-card | 4320; hours; mb |  | 1 | 0 | 1 | No validation | 0 |  | No validation studies found. Used as reference device in three studies | 0 |  |  | 2024-05-01 | 3.7 | 3.3
007_sia_wd_bit | Bittium | Bittium Faros 180 | https://www.bittium.com/medical/bittium-faros | 2019-01-01 | Current | Clinical | not allowed to share price details publicly | CPU + electrodes | Chest | 18 | L48xW13xH30 | 0 | 1; spot electrodes; 1 channel; aVF; 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 100; sternum | 0 | 0 | 0 | 0 | 1 | 168; hours | device charger | 90; minutes | 0 | 0 | 1 | HR pause; bradycardia; tachycardia |  | micro-USB; Blt classic |  | Bittium Faros Manager |  | mircoSD-card | 4320; hours; mb |  | 1 | 0 | 1 | No validation | 0 |  | No validation studies found. Used as reference device in one study | 0 |  |  | 2024-05-01 | 3.7 | 3.3
008_sia_wd_bit | Bittium | Bittium Faros 180L | https://www.bittium.com/medical/bittium-faros | 2022-05-01 | Current | Clinical | not allowed to share price details publicly | CPU + electrodes | Chest | 18 | L48xW13xH31 | 0 | 1; spot electrodes; 1 channel; aVF; 250 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 100; sternum | 0 | 0 | 0 | 0 | 1 | 336; hours | device charger | 90; minutes | 0 | 0 | 1 | HR pause; bradycardia; tachycardia |  | micro-USB; Blt classic |  | Bittium Faros Manager |  | mircoSD-card | 4320; hours; mb |  | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-05-01 | 4 | 3.3
009_sia_wd_our | Oura | Oura Ring Gen3 | https://ouraring.com/ | 2021-11-01 | Current | Consumer | 314; one-time purchase price including a device but not the required monthly app subscription (6 EUR) | Ring | Finger | 5 | sizes 4-15 | 1; transmission & reflection; green, red and infrared; 125; finger | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 50; finger | 0 | 0 | 1; 0.1; finger | 0 | 1 | 168; hours | device charger | 80; minutes | 0 | 1 | 0 | HR; Daily Activity Score; Daily Readiness Score; Daily Stress Score; Daily Sleep Score; Daily SpO2 averages; Session based HR; Session based HRV; Session based motion count; Sleep based HR; Sleep based HRV; Sleep based readiness; Steps | 5min: HR; 1day: activity, readiness, stress, sleep, SpO2; | BLE; WiFi | Windows NA; iOS 14 or higher; Android 8.0 or higher; macOS NA | Oura App; Oura on the web |  | built-in storage | 168; hours; mb | within the geographic region where data is collected | 1 | 0 | 1 | external | 4 | HR; RMSSD; AVNN; pNN50; SDNN; HF; LF; LF:HF ratio | High validity for HR and RMSSD and acceptable accuracy for nocturnal AVNN, pNN50, HF and SDNN in average-per-night tests with high error rates for LF and LF:HF ratio. | 1 | Participant compliance was high over several weeks. |  | 2024-09-24 | 4 | 7.3
010_sia_wd_min | MindWare Technologies | MindWare Mobile | https://www.mindwaretech.com/mindware-mobile/ | 2014-01-01 | Current | Research | 5499; electrodes and router comes at an extra price this is before tax (21%) | CPU + electrodes | Hip | 370 | L117xW79xH47 | 0 | 1; spot electrodes; 1 channel; 500 | 1; spot electrodes; 500; thorax | 1; 500; facial muscles | 1; spot electrodes; impedence-derived; 500; thorax | 1; spot electrodes; 500; thenar-hyperthenar | 0 | 0 | 1; 3 axes; 50; hip | 0 | 0 | 0 | 1; EOG | 0 | 48; hours | device charger | minutes | 0 | 0 | 1 | ACC; CO; DZ/dt; EMG; LVET; HR; PEP; PTT; Resp; RMSSD; RSA; SCL; Z0 | 1min: all | WiFi; SD-card |  | BioLab |  | SD-card | 48; hours; 4000; mb |  | 1 | 0 | 1 | No validation | 0 |  | No validation studies found. Used as reference device in one study | 0 |  |  | 2024-06-27 | 7 | 2.7
011_sia_wd_nov | Novacor | Diasys 3 plus | https://novacor.com/product/diasys-3-plus/ | 2020-01-01 | Current | Research | 2000; 1760 for the unit and 300 for the associated required software | CPU + strap | Hip | 195 | L85xW85xH25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; auscultatory and oscillometry; upper arm | 1; 3 axes; hip | 0 | 0 | 0 | 0 | 0 | hours | disposable batteries | 0; minutes | 0 | 1 | 0 | BP (mean); DBP; SBP; HR; PP | Manual/scheduled initiation of measurement | Cable USB-C | Windows NA; iOS NA; Android NA; macOS NA | HolterSoft Ultima ABPM | HolterSoft Ultima | built-in storage | hours; mb |  | 1 | 0 | 1 | external | 1 | DBP; SBP | The device is validated according to the AMI/ANSI/ISO 81060-2:2013 standard, ESH-IP revision 2010 and MEDDEV 2.7/1 rev 4 guidelines; and passes all criteria for both the DBP and SBP measures (i.e., for auscullatory and oscillometric measurements). | 0 |  |  | 2024-06-26 | 6.3 | 3
012_sia_wd_car | Carré Technologies Inc. | Hexoskin Proshirt | https://www.hexoskin.com/ | 2019-01-01 | Current | Research & Consumer | 600 | CPU + vest | Chest | 75 | Shirt: 2XS-4XL; CPU: L42xW72xH13 | 0 | 1; touch; 1 channel; lead I; 256 | 0 | 0 | 1; Respiratory Inductance Plethysmography; 2 belts; 128; rib cage and abdomen | 0 | 0 | 0 | 1; 3 axes; 64; near abdomen | 0 | 0 | 0 | 0 | 1 | 36; hours | device charger | 90; minutes | 0 | 1 | 1 | HR;RR;VM;VT;Steps;EE;Cadence;VO2max | 1min: HR;RR;VM;VT;Steps;EE;Cadence;VO2max | Own cable of the system (for transferring all collected data to the computer application); Bluetooth | Windows 7-8; iOS 7 or higher; Android 4.0.4 or higher; macOS 10.8 or higher | Hexoskin OneSync; Hexoskin Dashboard; HxConvertSourceFile | VivoSense is mentioned to work with Hexoskin data, in case the researchers would like a user-friendly interface to process the raw signal data | built-in storage | 600; hours; mb | USA | 1 | 0 | 1 | external | 12 | HR;RR;VT;VM;EE;Steps;HMI | Strong evidence is present that respiratory rate (RR) and heart rate (HR) can both be validly estimated with Hexoskin Proshirt during a variety of light, moderate, and intense activities, whether controlled or naturalistic. Only two studies one of which were naturalistic showed HR to not correspond with reference device values, and a study had a negative conclusion for RR. The reference devices used to compare the Hexoskin parameters were high-quality too. Overall, the accuracy of HR and RR at high intensity exercises decreased. Multiple studies showed test-retest reliability of HR and RR were high across 3-4 days. Most studies suggest tidal volume and minute volume (as well as inspiratory capacity and inspiratory reserve volume) also have acceptable validity, but this significantly decreases with more intense physical activity. Several studies found the tidal volume and minute volume to be out of alignment with a reference device. Energy expenditure (EE) was found to be invalid in all studies that assessed it. Findings for the accuracy of Steps were very mixed. While approximately half of studies that assessed it determined it is valid, other half of studies found it to be inaccurate. Hip Motion Intensity (HMI) was found to be sufficiently valid in one study. | 4 | Hexoskin is perceived by participants to be highly comfortable when used in a laboratory procedure. It can be used in those with chronic obstructive pulmonary disease when the use duration is several hours. Hexoskin can be preferred over wearable watches in studies involving individuals with dementia as the vest cannot be easily removed. Hexoskin use in an ambulatory study of a week, however, is problematic. Combined with an insufficient intrinsic motivation to wear and the need to repeatedly wet the electrocardiogram sensors, older individuals found it too burdensome to be worn in daily life for longer than a day. |  | 2023-11-28 | 7.7 | 3.7
013_sia_wd_app | Apple | Apple Watch SE | https://www.apple.com/apple-watch-se/ | 2020-09-01 | Current | Consumer | 240; iPhone is required | Watch | Wrist | 36 | L44xW38xH10.7 | 1; reflection; green and infrared; wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 1; 3 axes; wrist | 1 | 0 | 1; Microphone; Altimeter; Ambient Light Sensor | 1 | 18; hours | device charger | 90; minutes | 1 | 1 | 0 | HR;HRV-SDNN;Active Time;Steps;EE;Sleep Stages;Sleep Detection;Cadence;Distance;Menstrual Cycle Tracking;Sleep Latency; Sleep Efficiency | 1hour: HR, HRV-SDNN, steps, acitivyt time; 1day: HR, HRV-SDNN, Activity time, Distance, Steps, Sleep Detection, Sleep Stages, Sleep Latency | Blt 5.3; WiFi 4 | Windows NA; iOS iPhone 8 or later with iOS 16 or higher; Android NA; macOS NA | Apple Watch App |  | built-in storage; data has to be transmitted to the coupled Apple Health App (an iPhone is necessary) | hours; 32000; mb | Denmark for Europe, and 5 different server locations for the U.S.: Nevada, North Carolina, Arizona, Oregon, California | 1 | 1 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-07-09 | 2.3 | 4.2
014_sia_wd_app | Apple | Apple Watch Ultra 2 | https://www.apple.com/apple-watch-ultra-2/ | 2023-09-01 | Current | Consumer | 899; iPhone is required | Watch | Wrist | 61 | L49xW44xH14.4 | 1; reflection; green and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 1; 3 axes; wrist | 1 | 1; 0.01; wrist | 1; Barometer; Altimeter; Ambient light sensor; Water temperature;Depth gauge;Microphone | 1 | 36; hours | device charger | 90; minutes | 1 | 1 | 0 | HR;HRV-SDNN; Active Time;Steps;EE;Sleep Stages;Sleep Detection;Cadence;Distance;Menstrual Cycle Tracking;Sleep Latency; Sleep Efficiency; Afib detection; sinus rhythm detection | 1hour: HR, HRV-SDNN, steps, acitivyt time; 1day: HR, HRV-SDNN, Activity time, Distance, Steps, Sleep Detection, Sleep Stages, Sleep Latency | Blt 5.3; WiFi 3 | Windows NA; iOS iPhone Xs or later with iOS 17 or higher; Android NA; macOS NA | Apple Watch App |  | built-in storage; data has to be transmitted to the coupled Apple Health App (an iPhone is necessary) | hours; 64000; mb | Denmark for Europe, and 5 different server locations for the U.S.: Nevada, North Carolina, Arizona, Oregon, California | 1 | 1 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-03-24 | 2.3 | 4
015_sia_wd_app | Apple | Apple Watch Series 9 | https://www.apple.com/apple-watch-series-9/ | 2023-09-01 | Current | Consumer | 449; iPhone is required | Watch | Wrist | 32 | L41xW35xH10.7 | 1; reflection; green and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 1; 3 axes; wrist | 1 | 1; 0.01; wrist | 1; Barometer; Altimeter; Ambient light sensor;Microphone | 1 | 18; hours | device charger | 75; minutes | 1 | 1 | 0 | HR;HRV-SDNN; Active Time;Steps;EE;Sleep Stages;Sleep Detection;Cadence;Distance;Menstrual Cycle Tracking;Sleep Latency; Sleep Efficiency; Afib detection; sinus rhythm detection | 1hour: HR, HRV-SDNN, steps, acitivyt time; 1day: HR, HRV-SDNN, Activity time, Distance, Steps, Sleep Detection, Sleep Stages, Sleep Latency | Blt 5.3; WiFi 4 | Windows NA; iOS iPhone Xs or later with iOS 17 or higher; Android NA; macOS NA | Apple Watch App |  | built-in storage; data has to be transmitted to the coupled Apple Health App (an iPhone is necessary) | hours; 64000; mb | Denmark for Europe, and 5 different server locations for the U.S.: Nevada, North Carolina, Arizona, Oregon, California | 1 | 1 | 1 | No validation | 0 |  | No validity/reliability studies were conducted for this version. See Apple Watch Series 8. | 0 |  |  | 2024-03-24 | 3.3 | 5.3
016_sia_wd_app | Apple | Apple Watch Series 8 | https://support.apple.com/en-us/111848 | 2022-09-01 | Discontinued | Consumer | 400; iPhone is required | Watch | Wrist | 42 | L45xW38xH10.7 | 1; reflection; green and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 1; 3 axes; wrist | 1 | 1; 0.01; wrist | 1; Barometer;Altimeter; Ambient light sensor;Microphone | 1 | 18; hours | device charger | 75; minutes | 1 | 1 | 0 | HR;HRV-SDNN; Active Time;Steps;EE;Sleep Stages;Sleep Detection;Cadence;Distance;Menstrual Cycle Tracking;Sleep Latency; Sleep Efficiency; Afib detection; sinus rhythm detection | 1hour: HR, HRV-SDNN, steps, acitivyt time; 1day: HR, HRV-SDNN, Activity time, Distance, Steps, Sleep Detection, Sleep Stages, Sleep Latency | Bluetooth 5.3 and WiFi | Windows NA; iOS iPhone 8 or later with iOS 16 or later; Android NA; macOS NA | Apple Watch App |  | built-in storage; data has to be transmitted to the coupled Apple Health App (an iPhone is necessary) | hours; 32000; mb | Denmark for Europe, and 5 different server locations for the U.S.: Nevada, North Carolina, Arizona, Oregon, California | 1 | 1 | 1 | external | 4 | SpO2; Sleep time; Awake time; Sleep Stages; HR; Sleep Latency; Sleep Efficiency | Blood oxygen saturation (Sp02) was validated across a gold standard and had good accuracy, with a less than 4% deviation from true values. When two watches were worn at the same time by the same individual (same arm) at upper arm and wrist, heart rate measurements were in agreement, supporting device reliability. However, the test-retest reliability over time has not been investigated. The sleep duration and awake time durations had good validity as compared to polysomnography. The sleep stage detection validity is considered around acceptable (F1 score over 4 sleep stages was .49) though the Apple Watch 8 has been overperformed by all other wearables in a study that compared to polysomnography. Sleep latency was found to be very precise. Sleep efficiency also demonstrated acceptable to good validity. | 2 | Wearing the watch at the upper-arm by surgeons during surgery was shown to be feasible. A few participants who wore the watch and answered daily ecological prompts about the watch use states using less health/physiology oriented aspects of the watch (e.g., checking emails, listening to music) more often. They stated such functions do help with wearing the watch more. They did not see much point in wearing the watch while not working/exercising. However, these usability studies did not have sufficient sample sizes (1-4 people) and should be taken with a grain of salt. |  | 2024-08-27 | 3.7 | 6
017_sia_wd_bio | BIOPAC | Research Ring | https://www.biopac.com/product/research-ring/ | 2023-07-01 | Current | Research | 2020; one-time purchase price of a device, option to interface up to 9 rings with AcqKnowledge Ring software (€5284) | Ring | Finger | 9 | sizes 6-13 | 1; reflection; green; 400; finger | 1; Manually Initiated Recording; 1 channel; lead I; 800 | 0 | 0 | 0 | 1; gold plated dry electrodes; 400; finger | 0 | 0 | 0 | 0 | 0 | 1; 20; finger | 0 | 1 | 10; hours | device charger | 120; minutes | 0 | 1 | 1 | EDRindex, EDRamp, EDL, ΔEDL, HR, HRV, RSA | 1min: EDRindex, EDRamp, EDL, ΔEDL, HR, HRV, RSA | Bluetooth Low-Energy | Windows NA; iOS iOS 15.2 or higher; Android NA; macOS NA | Senstream App | AcqKnowledge Ring software from Biopac | stream and store data via the app - either locally on iPhone or sent to Senstream cloud; via AcqKnowledge - on your PC | 0; hours; 0; mb | USA | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-01-26 | 4 | 3.3
018_sia_wd_now | NOWATCH | NOWATCH | https://nowatch.com/ | 2023-01-01 | Current | Consumer | 450 | Watch | Wrist | 50 | D37xH10.9 | 1; reflection; green, red, infrared; 32; wrist | 0 | 0 | 0 | 0 | 1; dry electrode; wrist | 0 | 0 | 1; 32; wrist | 0 | 0 | 1; 1.67E-2; wrist | 1; Barometer | 1 | 144; hours | Device charger | minutes | 1 | 1 | 1 | HR; IBI; RR; activity count; activity type; cadence; sleep stages; steps | 1min: all | Bluetooth | Windows NA; iOS iOS 15.0 or higher; Android Android 10 or newer; macOS NA | NOWATCH app |  | built-in storage | 24; hours; mb | Europe | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-03-15 | 3.7 | 6
019_sia_wd_goo | Google | Pixel Watch 1 | https://store.google.com/gb/config/google_pixel_watch?hl=en-GB&selections=eyJwcm9kdWN0RmFtaWx5IjoiWjI5dloyeGxYM0JwZUdWc1gzZGhkR05vIn0%3D | 2022-10-01 | Discontinued | Consumer | 325 | Watch | Wrist | 36 | D41xH12.3 | 1; reflection; green, red, and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; wrist | 1; wrist | 1 | 0 | 1; Compass; Altimeter; Ambient Light Sensor | 1 | 24; hours | device charger | 110; minutes | 1 | 1 | 0 | HR; HRV; Breathing rate; Resting HR; Afib notifications; Daily Readiness score; Active Zone Minutes; Cardio Fitness Score; Exercise tracking; Stress Management score; Sleep score; Sleep time; Sleep stage detection | 1min: Blood oxygen saturation during sleep; 1hour: HR, Hourly activity, Stress management score, Resting HR; 1day: Sleep time, Sleep stage detection, RR during sleep, HRV, Skin temp, Steps | Bluetooth 5.0 | Windows NA; iOS NA; Android Android 8.0 or newer; macOS NA | Google Pixel Watch app, Fitbit app | Fitbit ECG app | built-in storage | hours; 32000; mb | several countries within the European Economic Area, United Kindom and Switzerland, controlled by Fitbit International Limited, an Irish company. | 1 | 1 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-03-15 | 2.7 | 4
020_sia_wd_goo | Google | Pixel Watch 2 | https://store.google.com/nl/product/pixel_watch_2?hl=nl | 2023-10-01 | Discontinued | Consumer | 399; Some features require Fitbit Premium (first 6 months free, after that 8.99EUR a month or 79.99EUR yearly) | Watch | Wrist | 31 | D41xH12.3 | 1; reflection; green, red, and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 1; dry electrode; wrist | 0 | 0 | 1; 3 axes; wrist | 1; wrist | 1 | 1; wrist | 1; Compass; Altimeter; Ambient Light Sensor; Barometer; Magnetometer | 1 | 24; hours | device charger | 75; minutes | 1 | 1 | 0 | HR; HRV; Breathing rate; Resting HR; Afib notifications; Daily Readiness score; Active Zone Minutes; Cardio Fitness Score; Exercise tracking; Stress Management score; Sleep score; Sleep time; Skin temp; SpO2; EDA; | 1min: Blood oxygen saturation during sleep; 1hour: Steps; HR (every 2 hours), Hourly activity, Stress management score, Resting HR; 1day: Sleep time, Sleep stage detection, RR during sleep, HRV, Skin temp, Daily blood oxygen saturation, Mean resting HR | Bluetooth 5.1 | Windows NA; iOS NA; Android Android 9.0 or newer; macOS NA | Google Pixel Watch app, Fitbit app | Fitbit ECG app, Fitbit EDA Scan app | built-in storage | hours; 32000; mb | several countries within the European Economic Area, United Kindom and Switzerland, controlled by Fitbit International Limited, an Irish company. | 1 | 1 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-03-15 | 4.3 | 6.3
021_sia_wd_cor | Corsano | Corsano Cardiowatch 287-2 | https://corsano.com/shop/ | 2022-10-01 | Current | Clinical | 1080; the price includes the watch, a blood pressure cuff, and platform access for 1 month - note that portal setup and configuration always comes with a one-off 2000 euro payment- subscription to the portal for a single bracelet over a year is 600 euros | Watch | Wrist | 19 | L25xW40xH10 | 1; reflection; green, red, and infrared; 32; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I; 256 | 0 | 0 | 0 | 1; dry electrode; wrist | 0 | 1; Cuffless Non-Invasive Blood Pressure algorithm; wrist | 1; 3 axes; 32; wrist | 0 | 0 | 1; 0.3; wrist | 0 | 1 | 264; hours | device charger | minutes | 0 | 0 | 1 | Activity; DBP; SBP; Core body temp; EDA; activity ECG; HRV; NIBP; HR; RR; Core body temp; Skin temp; Sleep; Spirometry; SpO2; Steps | 1min: all | Bluetooth BLE 5.0 | Windows NA; iOS NA; Android NA; macOS NA | Corsano Trials App, Medical cloud storage | Corsano Research Portal | built-in storage | 256; hours; mb | Corsano medical cloud storage in Switzerland or Germany | 1 | 1 | 1 | No validation | 0 |  | NA. See the synthesis for the previous version, Corsano Cardiowatch 287-1. | 0 |  |  | 2024-06-26 | 5 | 6.7
022_sia_wd_cor | Corsano | Corsano Cardiowatch 287-1 | https://corsano.com/shop/ | 2021-11-01 | Current | Clinical |  | Watch | Flexible | 15 | L36xW25xH9 | 1; reflection; green, red, and infrared; 32; wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 32; wrist | 0 | 0 | 0 | 0 | 1 | 168; hours | device charger | minutes | 0 | 0 | 1 | Activity; HR; HRV; RR; IBI; Afib; steps; cadence; speed; energy expenditure; calorid data; exercise type; sleep classifications | 1min: all | Bluetooth BLE | Windows NA; iOS NA; Android NA; macOS NA | Corsano Trials App, Medical cloud storage | Corsano Research Portal | built-in storage | 64; hours; mb | Corsano medical cloud storage in Switzerland or Germany | 1 | 1 | 1 | external | 1 | HR; IBI | PPG signal can determine HR and IBI with high accuracy in resting position (no speaking) in cardiovascular at-risk patient (also in the different subgroups for skin type, hair density, age, BMI and gender). | 0 |  |  | 2024-06-26 | 4.3 | 6
023_sia_wd_bio | Biospace Inc. | InBody BAND 2 | https://inbodyusa.com/products/band2/ | 2017-09-01 | Discontinued | Consumer | 149 | Watch | Wrist | 20 | L25.4xW17.78xH10.16 | 0 | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 0 | 0 | 0 | 1; Bioelectrical impedance | 1 | 336; hours | device charger | 60; minutes | 0 | 1 | 0 | %BF; Body Fat Mass; Skeletal Muscle Mass; HR; Steps; EE; Sleep Amount;Sleep Score;Sleep Stages;Active Time;Distance;Activity Detection | 1hour: Steps, EE, HR; 1day: Weight, %BF, Body Fat Mass, Skeletal Muscle Mass, Sleep amount, Duration of each sleep stage | Bluetooth | Windows NA; iOS 15.0 or higher; Android 10.0 or higher; macOS NA | InBody Wear App |  | built-in storage | 336; hours; mb |  | 1 | 0 | 1 | external | 1 | %BF | %BF (percentage body fat) can be measured both with excellent validity and reliability as measured in an adult population (mean age = 45). Other parameters have not been tested yet | 0 |  |  | 2023-12-03 | 2.7 | 3.7
024_sia_wd_fit | Fitbit | Sense | https://www.fitbit.com/global/us/products/smartwatches/sense | 2020-09-01 | Discontinued | Consumer | 150; Some features require a subscription app (first 6 months free, after that 8.99EUR a month or 79.99EUR yearly) | Watch | Wrist | 46 | L40.4xW40.4xH12.4 | 1; reflection; green, red, and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 1; Manually Iinitiated Recording; dry electrode; wrist | 0 | 0 | 1; 3 axes; wrist | 1; 3 axes; wrist | 1 | 1; wrist | 1; Ambient light; altimeter; | 1 | 144; hours | device charger | 120; minutes | 1 | 1 | 0 | HR; HRV; ns.SCR-MIR; Skin temp; Night SpO2; Night RR; Resting HR; Afib notifications; Daily Readiness score; Active Zone Minutes; Cardio Fitness Score; Steps; Distance; Calories; Exercise tracking; Wellness report; Stress Management score; Sleep log; Sleep Stages; Sleep Score | 1min: HR; 1hour: Distance, Calories, Hourly Activity, Exercise Count, Steps, Stress score, Skin temp; 1day: Night RR, HRV during night | Bluetooth 5.0; NFC | Windows NA; iOS iOS 15.0 or higher; Android Android 10.0 or higher; macOS NA | Fitbit App | Fitbit ECG app; Fitbit EDA Scan app | built-in storage | 168; hours; 4000; mb | several countries within the European Economic Area, United Kindom and Switzerland, controlled by Fitbit International Limited, an Irish company. | 1 | 1 | 1 | external | 5 | HR; SCL; Energy Expenditure (EG); Total sleep time (TST); Wake after sleep onset (WASO); sleep efficiency (SE) | Sufficient agreement for HR was reported by two studies under both dynamic and static conditions. However, one study found opposite results, with HR having proper agreement with a reference device for only two out of twelve participants. While skin conductance level was validly estimated, energy expenditure was not. Additionally, the device seems to overestimate Sleep Duration and Sleep Efficiency while underestimating Wake After Sleep Onset. | 0 |  |  | 2024-03-18 | 4.3 | 7.3
025_sia_wd_fit | Fitbit | Sense 2 | https://www.fitbit.com/global/us/products/smartwatches/sense2?sku=521BKGB | 2022-09-01 | Current | Consumer | 223; Some features require a subscription app (first 6 months free, after that 8.99EUR a month or 79.99EUR yearly) | Watch | Wrist | 38 | L40.5xW40.5xH12.3 | 1; reflection; green, red, and infrared; wrist | 1; Manually Initiated Recording; 1 Channel; Lead I | 0 | 0 | 0 | 1; dry electrode; wrist; palm | 0 | 0 | 1; 3 axes; wrist | 1; 3 axes; wrist | 1 | 1; wrist | 1; Ambient light; altimeter; | 1 | 144; hours | solar panel in device; Device charger | 120; minutes | 1 | 1 | 0 | HR; HRV; EDA (both scan-based and continuous); Skin temp; SpO2; Breathing rate; Resting HR; Afib notifications; Daily Readiness score; Active Zone Minutes; Cardio Fitness Score; Steps; Distance; Calories; Exercise tracking; Wellness report; Stress Management score; Sleep Score; Sleep Score Cumulative; | 1min: HR; 1hour: ns.SCR, Distance, Calories, Hourly Activity, Exercise Count, Steps, Stress score, Skin temp; 1day: Night RR, HRV during night | Bluetooth 5.0; NFC | Windows NA; iOS iOS 15.0 or higher; Android Android 10.0 or higher; macOS NA | Fitbit App | Fitbit ECG app; Fitbit EDA Scan app | built-in storage | 168; hours; 4000; mb | several countries within the European Economic Area, United Kindom and Switzerland, controlled by Fitbit International Limited, an Irish company. | 1 | 1 | 1 | No validation | 0 |  | No validation studies yet. See results for previous version (Fitbit Sense). | 0 |  |  | 2024-03-18 | 5.7 | 8
026_sia_wd_gar | Garmin | vivoactive 4 | https://www.garmin.com/en-US/p/643382 | 2019-09-01 | Discontinued | Consumer | 299 | Watch | Wrist | 50 | D40xH12.7 | 1; reflection; green, red, and infrared; wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 1; wrist | 0 | 0 | 1; Blood oxygen saturation sensor; Glonass; Galileo; Compass; Ambient light; Ambient temperature sensor; barometric altimeter | 1 | 192; hours | device charger | minutes | 1 | 1 | 0 | HR; HR zones; SpO2; VO2max; Active Time; Sleep duration; Steps; Stress score; Recovery time; Body Battery; Speed; Pace; Cadence; RR; Sleep stage detection (no rem) | 1min: HR; 5mins: Stress score (around 5 to 30mins per score, difficult to say as app shows data over 15 hours in a small graph) 1hour: Body battery, RR (day and night); 1day: Steps, Intensity minutes, Steps, Calories, Average resting HR, Sleep duration, Sleep stage duration (deep,light, awake, no rem); Notes: if an exercise session is manually started, the activity (including pace and cadence) is recorded per minute | Bluetooth, ANT+, Wi-Fi | Windows NA; iOS iOS 15.0 or higher; Android Android 7 or higher; macOS NA | Garmin Connect | Garmin Connect IQ Store; Garmin Dive; Garmin Golf; Garmin Jr.; Garmin ECG | built-in storage | 200; hours; 4000; mb | several countries, personal data regarding individuals who reside in a country in the European Economic Area, United Kindom or Switzerland is controlled by Garmin Würzburg GmbH; Data from customers in mainland China, is store in GarminChina Shanghai RHQ Co., Ltd's servers in mainland China; Others data will be stored on servers in United States of America, United Kingdom and/or Australia. | 1 | 1 | 1 | external | 4 | SDNN, HRVi, RMSSD, VLF, LF; Classification of inactive, light, moderate-to-vigorous physical activity and raw accelerometry signal (accelerometer); sleep duration | HRV measures - the one VRU study showed excellent accuracy for very low frequency HRV measures, but only moderate to inaccurate for high frequency metrics (especially in patients with cardio- or cerebrovascular diseases), in resting conditions (in the lab). Accelerometer - the validity and reliability of the raw accelerometer data shows neutral to positive results. The estimation of activity levels (inactive, light and moderate to vigorous physical acivity) appears promising, even though the results of raw accelerometer signal compared to a gold standards are moderate. Total sleep duration - the detected TSD showed poor correlation with the sleep diaries. | 0 |  |  | 2024-08-07 | 3.7 | 6.3
027_sia_wd_gar | Garmin | vivoactive 5 | https://www.garmin.com/en-US/p/1057989#specs | 2023-09-01 | Current | Consumer | 300 | Watch | Wrist | 26 | D42.2xH11.1 | 1; reflection; green, red, and infrared; wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; wrist | 0 | 1 | 0 | 1; Blood oxygen saturation sensor; Glonass/Galileo; Compass; Ambient light | 1 | 264; hours | device charger | 120; minutes | 1 | 1 | 0 | HR; Resting HR; HR zones; HRV; SpO2; VO2max; Active Time; Sleep duration; Stress score; Recovery time; Body Battery; Stress; Speed; Pace; Cadence; RR; Sleep recommendations; RR; Sleep stage detection | 1min: HR; 5mins: Stress score (around 5 to 30mins per score, difficult to say as app shows data over 15 hours in a small graph) 1hour: Body battery, RR (day and night), Distance; 1day: HRV (nighttime), Steps, Intensity minutes, Steps, Calories, Average resting HR, Sleep duration, Sleep stage duration (deep,light, awake, rem); Notes: if an exercise session is manually started, the activity (including speed and distance) is recorded per minute | Bluetooth, ANT+, Wi-Fi | Windows NA; iOS iOS 15.0 or higher; Android Android 7 or higher; macOS NA | Garmin Connect | Garmin Connect IQ Store; Garmin Dive; Garmin Golf; Garmin Jr.; Garmin ECG | built-in storage; data has to be transmitted to the coupled Garmin app. Data can be transmitted to your local computer storage | 264; hours; 4000; mb | several countries, personal data regarding individuals who reside in a country in the European Economic Area, United Kindom or Switzerland is controlled by Garmin Würzburg GmbH; Data from customers in mainland China, is store in GarminChina Shanghai RHQ Co., Ltd's servers in mainland China; Others data will be stored on servers in United States of America, United Kingdom and/or Australia. | 1 | 1 | 1 | external | 1 | GNNS | GNNS scores acceptably well on mountain terrain, in study with one participant. | 0 |  |  | 2026-01-07 | 3.3 | 6.8
028_sia_wd_who | WHOOP | WHOOP 4.0 | https://www.whoop.com/ | 2021-09-01 | Current | Consumer | 30; monthly subscription for a device | Adjustable strap | Wrist | 28 | L43.18XW27.9XH10.16 | 1; reflection; green, red, and infrared; flexible | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; flexible | 1; 3 axes; flexible | 0 | 1; flexible | 1; Pulse oximeter | 1 | 120; hours | Device charger; the device charger itself can be charged using a usb-c connector, and the device charger can then be clicked on the WHOOP itself for charging | 150; minutes | 1 | 1 | 0 | HR; Resting (sleep) HR; Blood oxygen saturation; Daily HRV-RMSSD during sleep; Monthly HRV-RMSSD during sleep; Daily RR during sleep; Averaged Monthly RR during sleep; Number of activities; Percentage daily strain; Activity type classification; Calories; Daily % Recovery; Sleep duration; Active time; Monthly maximal HR; Monthly averaged HR; Monthly Calories; Monthly averaged Sleep duration; Daily Bed Time; Averaged Monthly Bed Time; Daily Total Sleep Disturbances; Averaged Monthly Sleep Disturbances; Nightly Needed Sleep; Sleep Stages; Daily Skin Temperature; Stress score; Sleep performance | 1min: HR; 1hour: Stress score; 1day: Resting (sleep) HR, Blood oxygen saturation, Daily HRV-RMSSD during sleep, Daily RR during sleep, Number of activities, Percentage daily strain, Activity type classification, Calories, Daily % Recovery, Sleep duration, Active time, Daily Bed Time, Daily Total Sleep Disturbances, Nightly Needed Sleep, Sleep Stages, Daily Skin Temperature | Bluetooth 4.2 or higher; WiFi | Windows NA; iOS iOS 14 or higher, Android OS 10.0 or higher; Android iOS 14 or higher, Android OS 10.0 or higher; macOS NA | WHOOP App |  | built-in storage; the data should be synced to the app in mobile phone via Bluetooth | 336; hours; mb | for the European Economic Area, Switzerland, and UK, the data is stored in accordance with GDPR, exact server locations are not provided, the website notes for Taiwan residents, personal data may be stored in and transferred between any of the servers in the United States, Germany, Japan, United Kingdom, France, Canada, and India | 1 | 0 | 1 | No validation | 0 |  | No validation study of this version exists. See studies on the previous version of the device. | 0 |  |  | 2024-08-26 | 4 | 6.7
029_sia_wd_who | WHOOP | WHOOP 3.0 | https://shop.whoop.com/en-eu/collections/whoop3-0 | 2019-01-01 | Discontinued | Consumer | 30; monthly subscription for a device | Adjustable strap | Wrist | 18 | L43.18XW27.9XH10.17 | 1; reflection; green; flexible | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; flexible | 1; 3 axes; flexible | 0 | 0 | 0 | 1 | 120; hours | Device charger; the device charger itself can be charged using a usb-c connector, and the device charger can then be clicked on the WHOOP itself for charging | 150; minutes | 0 | 1 | 0 | HR; Resting (sleep) HR; Blood oxygen saturation; Daily HRV-RMSSD during sleep; Monthly HRV-RMSSD during sleep; Daily RR during sleep; Averaged Monthly RR during sleep; Number of activities; Percentage daily strain; Activity type classification; Calories; Daily % Recovery; Sleep duration; Active time; Monthly maximal HR; Monthly averaged HR; Monthly Calories; Monthly averaged Sleep duration; Daily Bed Time; Averaged Monthly Bed Time; Daily Total Sleep Disturbances; Averaged Monthly Sleep Disturbances; Nightly Needed Sleep; Sleep Stages; Daily Skin Temperature; Stress score; Sleep performance | 1min: HR; 1hour: Stress score; 1day: Resting (sleep) HR, Blood oxygen saturation, Daily HRV-RMSSD during sleep, Daily RR during sleep, Number of activities, Percentage daily strain, Activity type classification, Calories, Daily % Recovery, Sleep duration, Active time, Daily Bed Time, Daily Total Sleep Disturbances, Nightly Needed Sleep, Sleep Stages, Daily Skin Temperature | Bluetooth |  | WHOOP App |  | built-in storage; the data should be synced to the app in mobile phone via Bluetooth | 72; hours; mb | for the European Economic Area, Switzerland, and UK, the data is stored in accordance with GDPR, exact server locations are not provided, the website notes for Taiwan residents, personal data may be stored in and transferred between any of the servers in the United States, Germany, Japan, United Kingdom, France, Canada, and India | 1 | 0 | 1 | external | 4 | HR;HRV-RMSSD;Sleep time; Awake onset; Sleep onset; Sleep stage detection | Heart rate and heart rate variability (RMSSD) was shown to have good validity during sleep and physical activity that does not involve high gross body movement (e.g., dumbell curve). With exercise involving high gross body movement, the watch becomes misaligned with the reference device readings. Sleep detection as compared to polysomnography has high accuracy. Sleep stage detection, however, only moderately aligns with those stages identified by polysomnography. The heart rate and heart rate variability variations day-to-day are not any higher than what is expected for those values based on past literature, and it was argued that the device therefore does not add extra measurement related errors. | 1 | One usability study asking three individuals using WHOOP to reflect on their device use showed that individuals who are doing the sports that the device is marketed for (i.e. triathletes) are more likely to use the device. The monthly reports sent by the WHOOP can help users reflect on the impact of interventions/exercises they do on their body and recovery. |  | 2024-08-26 | 4 | 6.8
030_sia_wd_cir | Circular | Circular Slim | https://www.circular.xyz/ | 2023-12-01 | Discontinued | Consumer | 263 | Ring | Finger | 2 | sizes 4-16 | 1; reflection; green, red, and infrared; finger | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; finger | 0 | 0 | 1; finger | 1; Pulse oximeter | 1 | 120; hours | device charger | 45; minutes | 1 | 1 | 0 | HR; Morning Resting HR; RR during Sleep; Skin Temp; Stress Score; Energy Level; HRV-RMSSD during Sleep; Intermittent HRV-RMSSD; Steps; Sleep Duration; Sleep Disturbances; Sleep efficiency; Tranquility; Circadian Rhythm; REM sleep; Deep sleep; Time to fall asleep; Sleep debt; Awake times during sleep; Distance; EE;%maxHR; SpO2; Duration of sleep stages | 1min: HR, HRV, %maxHR, Activity intensity, Sp02; RR; 1hour: Energy level, Stress score; 1day: Sleep Score, Duration of sleep stages, Steps, Calories, Active minutes, EE | Bluetooth Smart | Windows NA; iOS iOS 11.0 or later; Android Android 6.0 or later; macOS NA | Circular |  | built-in storage; data has to be synced with the Circular App in smartphone at least every 10 days as the device cannot store for longer than 10 days | 240; hours; mb |  |  | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-03-29 | 2 | 4.5
031_sia_wd_mov | movisens | EcgMove4 | https://www.movisens.com/en/products/ecg-sensor/ | 2021-01-01 | Current | Research | 2089; Includes the CPU and one chest belt | CPU + strap | Chest | 26 | L11.5XW62.3XH38.6 | 0 | 1; 1 channel, Lead I; 1024 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 64; Chest | 1; 3 axes; 64; Chest | 0 | 1; 1; Chest | 1; Barometer | 1 | 72; hours | device charger | 120; minutes | 1 | 1 | 1 | HR; Step Count; RMSSD;EE;HRV;Movement Acceleration;Activity Class;Body Position; Skin Temp; Atmospheric Air Pressure; Angular Rate | 1min: all | USB cable; Bluetooth Smart 4.0 | Windows Windows 7, 8, 8.1, 10; iOS NA; Android NA; macOS NA | SensorManager | DataAnalyzer | built-in storage; the data is transferred from the local device storage to a PC using a USB cable | 336; hours; 4000; mn |  | 1 | 0 | 1 | external | 1 | HR; RMSSD; HF | All heart rate, heart rate variability in the form of RMSSD and in the form of HF (high frequency power) showed great alignment with the gold standard in both a laboratory procedure and in daily life, in an adult sample. Half-split reliability of the HR and HRV indices were high. Parameters derived from the accelerometer are not yet validated. The rate of valid data in ambulatory settings for HRV was higher than 90% and for daily life recordings higher than 80% (minimal amount of low quality data). | 0 |  |  | 2024-06-12 | 5.8 | 8
032_sia_wd_wit | Withings | Scanwatch | https://www.withings.com/eu/en/scanwatch-2 | 2021-11-01 | Current | Consumer | 299.95 | Watch | Wrist | 58 | D38xH13.2 | 1; Reflection; Green light; Wrist | 1; Manually initiated recording; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; Wrist | 0 | 0 | 0 | 0 | 1 | 720; hours | device charger | 120; minutes | 1 | 1 | 0 | Afib-MIR; HR; QRS; HR; QT; QTc; SpO2; activity tracking; sleep quality score; sleep stage light; sleep stage deep; sleep duration; sleep interruptions; sleep regularity; menstrual cycle tracking; Breathing disturbances; steps | 5min: HR; 1day: Sleep Quality Score, Sleep Stage Duration, Sleep SpO2; Daily SpO2, Breathing Disturbances, Activity tracking, Steps | Bluetooth BLE | Windows NA; iOS iOS 14 and higher; Android Android 10.0 and higher; macOS NA | Health Mate app |  | built-in storage | hours; mb | France | 1 | 1 | 1 | external | 10 | Afib; HR; QRS; PR; QT; QTc; SpO2; | HR can generally be accurately detected, with the best result in stationary conditions. AFib classification varies between studies, but has potential to be an accurate detection algorithm. Other ECG parameters such as the QRS, PR, and QTc show generally accurte results. However, one study used the device in different positions to obtain different ECG Leads which increase the amplitude, duration and morphology detections. The four studies evaluating the SpO2 show varying results; some very promising in multiple hypoxia condition, but the largest studies (n = 36-49) show less positive results and urge to interpret SpO2 results with care. | 0 |  |  | 2024-07-05 | 3 | 6.8
033_sia_wd_wit | Withings | Scanwatch 2 | https://www.withings.com/eu/en/scanwatch-2 | 2023-10-01 | Current | Consumer | 349.95 | Watch | Wrist | 35 | D38.4xH13.4 | 1; Reflection; Green, red and infrared light; Wrist | 1; Manually initiated recording; Lead I | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; Wrist | 0 | 0 | 1; Wrist | 1; Altimeter | 1 | 720; hours | device charger | 120; minutes | 1 | 1 | 0 | Afib-MIR; HR; Nighttime HRV; QRS; HR; QT; QTc; SpO2; activity tracking; sleep quality score; sleep stage light; sleep stage deep; sleep duration; sleep interruptions; sleep regularity; Skin temp; menstrual cycle tracking; Breathing disturbances; steps; RR | 5min: HR; 1day: Sleep Quality Score, Sleep Stage Duration, Sleep SpO2; Daily SpO2, Breathing Disturbances, Activity tracking, Steps, Nighttime HRV | Bluetooth BLE | Windows NA; iOS iOS 15 and higher; Android Android OS 9 and higher; macOS NA | Health Mate app |  | built-in storage | 168; hours; mb | France | 1 | 1 | 1 | No validation | 0 |  | NA (see previous version) | 0 |  |  | 2024-07-05 | 2.7 | 5.7
034_sia_wd_spa | Spacelabs | OnTrak | https://spacelabshealthcare.com/products/diagnostic-cardiology/abp-monitoring/ontrak/ | 2019-01-01 | Current | Research & Clinical | 3249 | CPU + strap | Hip | 172 | L10.15xH2.76xW7.0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; Oscillometry; Upper arm | 0 | 0 | 0 | 0 | 0 | 1 | 168; hours | Disposable batteries | 0; minutes | 1 | 1 | 0 | DBP;SBP;MAP;PP;HR | Manual/scheduled initiation of measurement | Cable | Windows Windows 7 SP1 Professional, Ultimate and Enterprise (64 bit), Windows 10 Professional and Enterprise (64 bit), Windows Server 2012 R2, Windows Server 2016 & 2019; iOS NA; Android NA; macOS NA | Sentinel |  |  | hours; mb |  | 1 | 1 | 1 | external | 1 | SBP; DBP | The Spacelabs 90227 OnTrak device fulfils the requirements of three recognized validation protocols (European Hypertension Society International Protocol revision 2010 (ESH-IP2), the British Hypertension Society Protocol (BHS) , and the American National Standards Institute/Association for the Advancement of Medical Instrumentation/International Standards Organization (ANSI/AAMI/ISO) 81060-2:2009 protocol) in adults and children and can therefore be recommended for clinical use in standard function. | 0 |  |  | 2024-07-12 | 6.7 | 2.3
035_sia_wd_gar | Garmin | Vivosmart 5 | https://www.garmin.com/en-US/p/782585/ | 2022-04-01 | Current | Consumer | 150 | Watch | Wrist | 25 | L19.5xW10.7xH217 | 1; Greend, red and infrared light; Wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; Wrist | 0 | 0 | 0 | 1; Ambient light sensor; blood oxygen saturation sensor | 1 | 168 | device charger | 90 | 1 | 1 | 0 | RR; Sleep score; Blood oxygen during the day; Blood oxygen during sleep; HR;Daily resting pulse rate; Fitness age; Body battery level; Daily stress score;Steps;EE;Distance;Intensity minutes;HR zones; HR alerts; % HR max; HR calories; Duration of each sleeping stage; Incident Detection; Sleep stage classification (light, deep, REM, awake) | 1min: HR, Stress Score, Body Battery; 1hour: RR; 5hours: Blood oxygen saturation; 1day: Sleep score, Daily stress score, Duration of each sleeping stage, Fitness age | Bluetooth Smart; ANT+ | Windows NA; iOS Apple iOS 16.0 and above; Android Android 7.0 OS and above; macOS NA | Garmin Connect App | Garmin Express | built-in storage | 339; hours; mb | data of individuals who reside in the European Economic Area, United Kindom or Switzerland is stored by Garmin Würzburg GmbH; Data from customers in China, is store in GarminChina Shanghai RHQ Co., Ltd's servers in mainland China; Others data is stored on servers in United States of America, United Kingdom and/or Australia. | 1 | 0 | 1 | No validation | 0 |  | The specific version of this device has no validity/reliability studies. See Garmin Vivosmart 4. | 0 |  |  | 2024-08-26 | 3.3 | 7
036_sia_wd_gar | Garmin | Vivosmart 4 | https://www.garmin.com/en-US/p/605739/ | 2018-09-01 | Discontinued | Consumer | 150 | Watch | Wrist | 17 | L15xW10.5xH197 | 1; Greend, red and infrared light; Wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; Wrist | 0 | 0 | 0 | 1; Ambient light sensor; blood oxygen saturation sensor; barometric altimeter | 1 | 168; hours | device charger | 129; minutes | 1 | 1 | 0 | Sleep score; Blood oxygen during the day; Blood oxygen during sleep; HR;Daily resting pulse rate; Fitness age; Body battery level; Daily stress score;Steps;EE;Distance;Intensity minutes;HR zones; HR alerts; % HR max; HR calories; Duration of each sleeping stage; Incident Detection; VO2max; Sleep stage classification (light, deep, REM, awake) | 1min: HR, Body battery, Stress Score; 5min: SpO2%; 1day: Sleep score, Daily stress score, Duration of each sleeping stage, Fitness age | Bluetooth | Windows NA; iOS Apple iOS 16.0 and above; Android Android 7.0 OS and above; macOS NA | Garmin Connect App | Garmin Express | built-in storage | 339; hours; mb | data of individuals who reside in the European Economic Area, United Kindom or Switzerland is stored by Garmin Würzburg GmbH; Data from customers in China, is store in GarminChina Shanghai RHQ Co., Ltd's servers in mainland China; Others data is stored on servers in United States of America, United Kingdom and/or Australia. | 1 | 0 | 1 | external | 4 | HR; sleep estimates (Total sleep time (TST), time in bed (TIB), wake, wake after sleep onset (WASO), sleep onset latency (SOL), sleep stages, REM latency, sleep onset and sleep end); step count; SpO2 | Garmin Vivosmart 4 has shown acceptable agreement with a chest monitor for HR among individuals with severe disabilities participating in the indirect movement program. It was also shown to be sufficiently acurate for estimating sleep duration but not for sleep architecture, for example detecting sleep stages. Intra-device reliability for sleep estimates was assessed in one participant with two devices over 23 nights and was reported as high. While the watch is valid at counting steps at regular walking speeds, it has compromised validity at slow walking speeds among healthy adults and no validity for step counting among Parkinson's patients. The device was also tested with Chronic Obstructive Pulmonary Disease patients on SpO2 estimation showing insufficient validity due to overstimation of SpO2 at low- and underestimation at high saturation levels. | 1 | Garmin Vivosmart (previous version) was one of the devices evaluated in a user acceptence study with non-metrapolitan cancer survivors who wore it for two weeks. The device was seen positively due to its aesthetics, comfort of wear and accessible design with bigger display and font being well perceived by older users. |  | 2025-02-17 | 3.7 | 7.5
037_sia_wd_gar | Garmin | Index BPM Smart Blood Pressure | https://www.garmin.com/en-US/p/716808/ | 2022-09-01 | Current | Consumer | 176 | CPU + strap | Upper arm | 320 | L58xW145xH47 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; Manometer; Upper arm | 0 | 0 | 0 | 0 | 0 | 1 | 6480; hours | Disposable batteries | 0; minutes | 1 | 1 | 0 | DPB;SBP;HR | Manual/scheduled initiation of measurement | Wi-Fi; Bluetooth | Windows NA; iOS NA; Android NA; macOS NA |  | Garmin Connect App | 100 measurements | 100; measurements | data of individuals who reside in the European Economic Area, United Kindom or Switzerland is stored by Garmin Würzburg GmbH; Data from customers in China, is store in GarminChina Shanghai RHQ Co., Ltd's servers in mainland China; Others data is stored on servers in United States of America, United Kingdom and/or Australia. | 1 | 1 | 1 | internal | 1 | SBP;DBP;HR | The internal clinical validation study is not published in a peer-reviewed journal but it helped receive an FDA clearance. In this study, the device can measure both blood pressure (+/- 2%) and pulse rate (+/- 5%) validly as revealed by measurements within the error margins determined by the American Heart Association. As study sample characteristics and task details were not given, neutral verdicts were given for blood pressure and pulse rate monitoring. | 0 |  |  | 2024-08-26 | 4.2 | 2
038_sia_wd_sky | SkyLabs | CART BP | https://www.skylabs.io/en | 2023-03-01 | Current | Clinical & Consumer | 447 | Ring | Finger | 4 |  | 1; Reflection; Green, red and infrared light; 175; Finger | 0 | 0 | 0 | 0 | 0 | 0 | 1; Machine learning model using PPG data; Finger | 0 | 0 | 0 | 0 | 0 | 1 | 48; hours | device charger | 120; minutes | 0 | 0 | 0 | HR; HRV; DBP; SBP; BPV | 5min: HR, DBP, SBP, BPV | Bluetooth | Windows NA; iOS iOS 10.0 and later; Android Android 6.0 and later; macOS NA | CART BP pro |  | built-in storage | 24; hours; mb |  | 0 | 0 | 0 | No validation | 0 |  |  | 0 |  |  | 2024-09-04 | 1.7 | 4
039_sia_wd_ama | Amazfit | Helio Ring | https://us.amazfit.com/products/amazfit-helio-ring | 2024-05-01 | Current | Consumer | 300 | Ring | Finger | 4 | sizes 8, 10, 12 | 1; Reflection; dual-color LED (specific lights types not provided); Finger | 0 | 0 | 0 | 0 | 1; Dry electrode; Finger | 0 | 0 | 1; 3 axes; Finger | 1; 3 axes; Finger | 0 | 1; Finger | 1; Blood oxygen saturation sensor | 1 | 48; hours | device charger | 120; minutes | 1 | 1 | 0 | Readiness score; HR; Blood oxygen saturation; Stress level; Sleep stages; Sleep breathing quality; Daytime naps; Sleep score; HR recovery; EDA; HRV; Emotional response detection; Active minutes; Steps; EE; Physical recovery; Mental recovery; Breathing quality; VO2max; Training load; Nocturnal resting HR | 1min: Speed, Cadence;5min: HR, Altitude;1day: Sleep score, Nocturnal resting HR, Breathing quality, Skin temp, Sleep stage durations (light, deep, REM), Count of awake times during sleep, Daily HR, VO2max, Training load | BLE | Windows NA; iOS iOS 14.0 and above ; Android Android 7.0 and above; macOS NA | Zepp App |  | built-in storage | hours; mb | the server country is not specified, but the company is based in China | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-09-04 | 3 | 5.2
040_sia_wd_bio | BioStrap | Biostrap Kairos | https://shop.biostrap.com/products/biostrap-kairos-pilot-set | 2023-07-01 | Current | Clinical & Consumer | 275 | Adjustable strap | Flexible | 10 | L36xW22xH12.5 | 1; Reflection; two green and 1 infrared light; Wrist | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; Wrist | 1; 3 axes; Wrist | 0 | 0 | 1; Blood oxygen saturation sensor | 1 | 120; hours | device charger | minutes | 0 | 1 | 0 | HR; HRV; RR; SpO2; sleep analysis; steps; activity metrics; arterial and peripheral elasticity | 5min: HR, HRV, SpO2; sleep analysis; steps; activity metrics | Bluetooth 5.0 | Windows NA; iOS NA; Android NA; macOS NA | Vital Science application | Remote physiological monitoring web dashboard | built-in storage | hours; 64; mb |  | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-09-04 | 2.3 | 4.3
041_sia_wd_car | CardioMood | CardioMood 287-2 | https://cardiomood.com/product/cardiowatch-287-2 | 2023-01-01 | Current | Research & Consumer | 499 | Wrist strap | Wrist |  | 40Lx25Hx10W | 1; Green, red and infrared light; 64-128; Wrist | 0 | 0 | 0 | 0 | 1; Dry electrodes; Wrist | 0 | 0 | 1; 3 axes; 64-128; Wrist | 0 | 0 | 0 | 1; Heat flux; Blood oxygen saturation sensor | 1 | 240; hours | device charger | minutes | 0 | 1 | 1 | Steps; Distance; EE; Core Body temp; HRV-MRR; HRV-SDNN; HRV-RMSSD; HRV-PNN50;HRV-MO; HRV-AMO; HRV-SI; IBI; HR; SpO2; Sleep stage classification (light, deep, awake, REM); RR; Performance; Sleep Tranquility score; Sleep Consistency Score; Sleeptime; Resting HR; Readiness score; EDA; Stress | 1mins: HR and HRV indices IF Preventicus heartbeats software is used; 5mins: Readiness score; 10mins: HR; 1hour: EE (calories), steps; 1day: Core body temperature, HRV, RR | Bluetooth 5.0 | Windows NA; iOS iOS 12.1 or higher; Android Android 8.0 or higher; macOS NA | CardioMood Trends | Preventicus heartbeats | built-in storage | hours; 64; mb | Preventicus heartbeats | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2025-02-12 | 2.3 | 4.8
042_sia_wd_sam | Samsung | Galaxy Ring | https://www.samsung.com/us/rings/galaxy-ring/ | 2024-08-01 | Current | Consumer | 449 | Ring | Finger | 3 | sizes 5-15 | 1; Reflection; Green light; Finger | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; Finger | 0 | 0 | 1; Finger | 1; Blood oxygen saturation sensor | 1 | 168; hours | device charger | 80; minutes | 0 | 1 | 0 | HR;SpO2;Sleep duration;Energy score; Workout tracking; Skin Temp; Active sleep score; Snoring analysis; Movement during sleep; Sleep latency; RR; Sleep quality score; Sleep routine score; Min HR; Max HR; Period onset estimation using skin temp | 1hour: SpO2 during sleep, Skin temp, Min HR, Max HR, Movement during sleep;1day: Sleep duration, Sleep stage durations (light, deep, REM, awake), Daily HR, Restfulness score, Mental recovery score, Sleep Score, Physical recovery | BLE 5.4 | Windows NA; iOS NA; Android Android 11 or above; macOS NA | Samsung Health App |  | built-in storage | hours; 8000; mb | servers are located in South Korea, U.S.A., several locations in Europe, among other locations | 1 | 0 | 0 | No validation | 0 |  |  | 0 |  |  | 2024-09-04
043_sia_wd_goo | Google | Pixel Watch 3 | https://store.google.com/product/pixel_watch_3?hl=nl | 2024-01-01 | Current | Consumer | 399 | Watch | Wrist | 31 | D41xH12.3 | 1; Reflection; Green, red and infrared light; Wrist | 1; Manually initiated recording; one channel similar to Lead I | 0 | 0 | 0 | 1; Dry electrodes; Wrist | 0 | 0 | 1; 3 axes; Wrist | 1; Wrist | 1 | 1; Wrist | 1; SpO2; Compass; Altimeter; Ambient Light Sensor; Barometer; Magnetometer | 1 | 36; hours | device charger | 60; minutes | 1 | 1 | 0 | Steps; Distance; Floors; EE; Readiness score; HR; Sleep quality; Cardio load; Target cardio load; HRV; Resting HR; Sleep stages (light, deep, awake REM); Sp02%; Skin temp variation; RR; Stress management score; Physical calmness; Activity balance; | 1min: HR; 1hour: Steps, Distance, EE; 1day: Resting HR, RR, SpO2%, Skin temp variation | Bluetooth 5.0; Ultra-wideband; Wi-Fi 802.11 | Windows NA; iOS NA; Android Android 10.0 or higher; macOS NA | Google Pixel Watch; Fitbi App; Fitbit ECG App |  | built-in storage | hours; 32000; mb |  | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2025-02-13
044_sia_wd_adi | AD Instruments | Equivital eq02+ Lifemonitor | https://equivital.com/products/eq02-lifemonitor | 2021-01-01 | Current | Research | 1977 | CPU + vest | Chest | 38 | 78x55x11 | 0 | 1; Two channels similar to lead I and lead II, based on touch connectivity (vest to skin); 256; Chest | 0 | 0 | 1; Using Strain Gauge Transducer; 25.6; Chest | 0 | 0 | 1; Optional using cuff-based compatibble ancillary sensor (Note, not standard!); Upper arm | 1; 3 axes; 25.6-256; Chest | 0 | 0 | 1; Accuracy +/- 0,1 degree Celsius; measures every 15 seconds; 0.067 | 1; Core temperature; Oximeter | 1 | 48; hours | Device charger; either a dedicated USB compatible cable, or a docking station for charging up to 6 modules | 60; minutes | 1 | 1 | 1 | HR; HR confidence index; IBI; RR; RR confidence index; Skin Temp; Physiological Welfare Index state; Fall state | 15sec: all, 5sec:HR (if configured) | cable; system specific-USB cable; Bleutooth Compliance version 2.1 | Windows NA; iOS NA; Android NA; macOS NA | Equivital™ Manager | Equivital™ Qiosk; eqView Professional (Non-Diagnostic); Equivital™ Black Ghost eqView Mobile; | built-in storage | 1200; hours; 8000; mb |  | 1 | 1 | 1 | external | 4 | HR; RMSSD; SDNN; pNN50; average of NN intervals; Estimated Core Body Temp; Simulated breath hold, simulated chest motion signal linearity | Valid and reliable for estimating core body temperature with a slight overestimation of 0.1°C. It provides reliable HR, RMSSD and pNN50 measurements but only moderate reliability for SDNN and other HRV parameters , with artifacts affecting accuracy. It's respiratory sensor is not accurate for estimating chest movement or breath hold as shown by a simulated benchtop test. | 0 |  |  | 2022-02-05 | 6.8 | 3
045_sia_wd_med | Medtronic Zephyr | BioHarness 3.0 | https://www.zephyranywhere.com/system/overview | 1899-12-31 | Current | Research & Consumer | 766; Including biomodule, one charger and a strap; optionally add GPS for 190 and GPS charger 5 bay for 340, gateway for live monitoring 450; Omnisense software 4530, and SDK kit for downloading and streaming and matlab script 170; all excluding VAT | CPU + electrodes | Chest | 18 | 28Dx7H | 0 | 1; One channel similar to lead I; 256 | 0 | 0 | 1; Measured via belt at rib cage; 18; Thorax | 0 | 0 | 0 | 1; 3 axes; 100; Thorax | 0 | 1; Optional addition can be purchased | 0 | 0 | 1 | 35; hours | device charger | 180; minutes | 0 | 0 | 0 | ACC; RR; HR; HRV; % HR max; % HR at anaerobic threshold; SpO2; posture detection; estimated core temp; impact; activity; caloric burn; activity metrics | 1min: all | cable; system specific-USB cable; Bleutooth Compliance version 2.1 + EDR and version 4.0 | Windows Windows 8 or higher; iOS NA; Android NA; macOS NA | SDK Kit | Omnisense software | built-in storage | 450; hours; mb |  | 1 | 1 | 1 | external | 6 | HR; HRV; RR; accelerometry; step counts | The Zephyr BioHarness 3.0 can reliably measure HR and RR with high test-retest reliability results (72 hours apart). The convergent validity of HR and HRV is examined in merely one individual and shows positive results for a resting condition for HR, whereas HRV showed poor results. The convergent validity of RR is also assessed in merely one individual but showed positive results. The accelerometer signal for step counts shows valid accuracy when compared to video recording; with the most accurate results for jogging, a slight underestimation for walking and small overestimation for running. | 0 |  |  | 2024-08-14 | 6.5 | 5.4
046_sia_wd_rin | RingConn | RingConn Gen2 | https://ringconn.com/products/ringconn-gen-2?variant=48162846769460 | 2024-08-01 | Current | Consumer | 275 | Ring | Finger | 3 | sizes 6-14 | 1; Green, red and infrared; Finger | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; finger | 0 | 0 | 1; finger | 1; SpO2; finger | 1 | 288; hours | device charger | 120; minutes | 0 | 0 | 0 | Activity intensity, HR, HRV, RR, SpO2, Skin Temperature, Slepp Stages, Sleep Apnea Detection, Steps, Stress Index | 1min: SpO2, 5min: HR, 1 hour: activity intensity, sleep stages, steps | Bluetooth 5.0; Low power bluetooth module | Windows NA; iOS NA; Android NA; macOS NA | RingConn App | Apple Health, Google Fit or others to keep your data for longer | built-in storage | 168; hours; mb | United Kingdom for users in the United States, the United Kingdom and the European Economic Area (EEA), and servers in People's Republic of China for users in People's Republic of China. | 1 | 0 | 1 | internal | 0 | OSAHS detection model (using SpO2 and PPG signals) | Based on the previous version, i.e., the RingConn Gen1: Model appears accuracte and specific to detect OSAHS using the RingConn signals. | 0 |  |  | 2024-10-31
047_sia_wd_ult | Ultrahuman | Ultrahuman Ring AIR | https://www.ultrahuman.com/ | 2023-01-01 | Current | Consumer | 379 | Ring | Finger | 2 | sizes 5-14 | 1; reflection; green, red and infrared; finger | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; finger | 1; 3 axes; finger | 0 | 1; finger | 1; SpO2; finger | 1 | 144; hours | device charger | 90; minutes | 0 | 1 | 0 | HR;Sleep stages (with REM); Sleep efficiency score; Sleep restfulness score; Sleep time; HRV-RMSSD; HRV-SDNN; Stress score; Step Count; Distance; EE; VO2max; Atrial Fibrillation detection; Cardio Adaptability detection; Circadian Alignment Suggestions; RR; Blood oxygen saturation; Dynamic recovery score; Sleep movements; Movement score; Active minutes; Workout frequency; Detection of Stress periods; Detection of Relaxed periods; Detection of Stimulated periods; Detection of Active periods; Resting heart rate; Skin temp | 5min: HRV; 30min: HR, Active minutes, Stress score; 1hour: Sleep movements; 1day: Sleep stages (with REM); Sleep efficiency score; Sleep restfulness score; Sleep time | Bluethooth low energy (BLE5) | Windows NA; iOS iOS 15 or later; Android Android 6 or later; macOS NA | Ultrahuman App |  | built-in storage | 168; hours; mb | U.S., possibly india | 0 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2024-11-03 | 2 | 4.3
048_sia_wd_nua | Nuanic | Nuanic Ring | https://nuanic.com/ | 2023-04-01 | Current | Research | 650; price depends on the number of rings ordered; mobile app, access to raw data, 12 months of complimentary Nuanic Cloud and dedicated technical support is included | Ring | Finger | 7 | sizes 7, 9, 10, 12 | 0 | 0 | 0 | 0 | 0 | 1; dry electrode; 3-16; finger | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 144; hours | device charger | 30; minutes | 0 | 0 | 1 | DNE (Propriatory Nuanic index (1-100) derived from EDA and accelerometry) | 1min: DNE | Bluetooth LE | Windows NA; iOS iOS 14 or higher; Android Android 5 or higher; macOS NA | Nuanic App |  | built-in storage | 72; hours; mb | Nuanic Could will become available and will be GDPR compliant with data stored within the European Economic Area. Data might me processed outside of this area. Could storage will be optional and the integrated storage can still be used. | 1 | 0 | 1 | External | 0 | nSCR;AmpSum;SCR;ISCR;PhasicMax; | Only one study was identified which examined the validity of the initial prototype of the Moodmetric Ring (the first version of Nuanic Ring). Results show sufficient validity of the EDA recording, however issues with data transfer were encountered which the authors attributed to the prototype form of the ring. It is also unclear how and if the Moodmetric Ring differed from the current Nuanic Ring so the validity conclusions should be taken with caution. | 0 |  |  | 2024-11-06 | 3 | 4.7
049_sia_wd_cos | Cosinuss | c-med alpha | https://www.cosinuss.com/en/ |  | Current | Clinical |  | Earable | Ear | 7 | L55.2xH58.6xH10.0 | 1; reflection, read and infrared; 200; ear | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1; 3 axes; 300; ear | 0 | 0 | 0 | 1; SO2, Core Temperature | 1 | 12; hours | device charger | 60; minutes | 0 | 0 | 0 | Core Temperature; HR; RR; SpO2 | 1min: all | Bluetooth BLE 5.0 | Windows NA; iOS NA; Android NA; macOS NA | c-med° App, and Cosinuss Health App0 | Cosinuss Lab system |  | hours; mb |  | 1 | 0 | 1 | External | 2 | Core temperature, HR and SpO2 | The validation studies regarding SpO2 and HR show positive results and demonstrate the c-med alpha can accurately measure these parameters in resting condition (according with ISO 80601-2-61 standard). Note this was in internal study. The external study assessing Core Temperature (Tc) estimation showed poor results and conclude Tc cannot be estimated using the c-med alpha. | 1 | Using the c-med alpha showed to be feasible for low-intensity care settings, such as general wards. Patient and clinical staff were generaly positive about the sensor acceptance and ease of use; the support regarding system use was experience as good. The system set up was fast (within 5-10min). |  | 2025-02-04 | 6.3 | 4.3
050_sia_wd_lum | Lumia Health | Lumia | https://lumiahealth.com/ | 2025-01-01 | Upcoming | Consumer | 379; plus a subscription cost of 16.58 per month for a yearly subscription | Earable | Ear | 1 |  | 1; reflection, infrared at the posterior auricular artery; ear | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 48; hours | solar panel in device, device charger | 240; minutes | 0 | 1 | 0 | Blood flow to the head; HR; HRV; Daily average flow index; Daily minimum flow index | 5min: all | Bluetooth 5.0 | Windows NA; iOS iOS 16 or later; Android NA; macOS NA | Lumia: Manage your dizziness app (available in US only) |  | built-in storage | hours; mb |  | 0 | 0 | 0 | No validation | 0 |  |  | 0 |  |  | 2024-02-13
051_sia_wd_plu | Plux | BioSignal kit | https://www.pluxbiosignals.com/products/4-channel-biosignals-kit | 2024-04-01 | Current | Research | 2510; 4-Channel biosignalsplux Kit costs 2359.50euro when bought separately. When combined in a kit builder set, the 4-channel kit costs 1950.00euro, where sensors with varying prices can be added; an example is given using ECG, EMC, EDA and accelerometer sensors. Extra costs for add-on software packages are applicable (e.g., 598.95euro) | CPU + electrodes | Flexible | 45 | 54Hx85Lx10W | 1; transmission; 500; finger | 1; electrodes as contact type; 100 | 0 | 1; 25-500; flexible | 1; inductive respiration (RIP) or piezoelectric respiration belt options; belt at chest | 1; adhesive electrodes; index finger and palmar side of hand | 1; 1 channel; 0.8-48.23 | 0 | 1; 3 axes; 0-50 | 0 | 0 | 1; flexible | 1; Blood pulse volume finger clip sensor; SpO2 finger clip sensor; Force sensor; Goniometer; Electrooculography sensor;Light (LUX) sensor; Functional Near-Infrared Spectroscopy (FNIRS) sensor; Piezo-Electric Respiration; Electrogastrography (EGG) sensor. |  | 12; hours | device charger | 150; minutes | 0 | 0 | 1 | Breathing cycle amplitude parameters (max, mean, SD, root-mean-square and total energy estimation); breathing frequency parameters; RR; Breath-to-Breath Interval (BBI) parameters (min, max, mean, SD root-mean-square and total BBI differences); HRV parameters (SDNN, RMSSD, NN20, pNN20, pNN50); mean HR; NN50; ECG frequency parameters (VLF, LF, HF, peak frequency, absolute powers, normalized powers, LF/HF ratio); EMG activation-specific parameters (start, end, duration of muscle activation, latency, peak-to-peak amplitudes, root-mean-square, integral over activation segment); electrodermal response (EDR) latency; EDR amplitude, EDR rise time; EDR recovery time (to 50% and to 63% of EDR amplitude); Response peak after stimulus, total power estimation; maximum foce (N); force reproduction rate; jump height; fleight time during jump | 1min: all | Bluetooth 2.0+EDR, USB calbe to download stored data from memory (optional extra) | Windows NA; iOS NA; Android NA; macOS NA | OepnSignals (r)evolution |  | built-in storage | 111; hours; 8000; mb |  | 1 | 0 | 1 | No validation | 0 |  |  | 0 |  |  | 2025-02-03 | 7 | 2.5

## 001_sia_wd_emp
device_id | study_citation | year | journal | external_internal | population | sample_size | age | reference_device | reference_device_other | events | events_included | param_studied | test_retest_time | conclusion_authors | sum_val_rel | sum_usability | final_verdict
001_sia_wd_emp | Koskimäki, H., Monttinen, H., Siirtola, P., Huttunen, H. L., Halonen, R., Röning, J., & Acm. (2017). Early Detection of Migraine Attacks Based on Wearable Sensors - Experiences of Data Collection using Empatica E4. Proceedings of the 2017 ACM International Joint Conference on Pervasive and Ubiquitous Computing and Proceedings of the 2017 ACM International Symposium on Wearable Computers. | 2017 | Proceedings of the 2017 ACM International Joint Conference on Pervasive and Ubiquitous Computing and Proceedings of the 2017 ACM International Symposium on Wearable Computers | External | Clinical: migraine patients | 11 | [5 empty] | Survey on user experience (what was good about the device and what caused problems) |  | "From the perspective of the data collection the most interesting aspect was that all the users would volunteer again to the study. Of course, when taken into account that only one of the answerers had the device in use more than two months it can be biased that way that the ones actually having problems with the device did not want to take part in the survey. In actual usage of the device the answerers either reported having no trouble, slight trouble or moderate trouble while the most positive feedback was given to the easiness of first time use and to the way that data were uploaded to cloud and how to recharge the device. Nevertheless, the daily upload and charging were the most negative aspects of the device."" |  | Most participants found the device easy to use and appreciated the data download and charging functionality. At the same time however, charging and data handling were reported as most problematic aspects of the device. | Positive for: ease of use; Neutral for: charging the device; downloading of data
001_sia_wd_emp | Pietilä, J., Mehrang, S., Tolonen, J., Helander, E., Jimison, H., Pavel, M., &amp; Korhonen, I. (2017). Evaluation of the accuracy and reliability for photoplethysmography based heart rate and beat-to-beat detection during daily activities. IFMBE Proceedings, 145–148. https://doi.org/10.1007/978-981-10-5122-7_37 | 2018 | IFMBE Proceedings | Ilkka Korhonen is a co-founder of PulseOn Technologies Ltd, Espoo, Finland. The authors declare that they have no other conflict of interest. | Healthy: young males | 25, 20 after exclusion | M = 27.7 SD = 3.5 | Firstbeat Bodyguard 2 |  | 5 | Sitting Standing Household work Mental arithmetic Cycling | HR; IBI; RMSSD; |  | "The results of this study suggest that the PPG wrist-worn monitors may have an acceptable accuracy and reliability in terms of HR and beat-to-beat detection during different daily activities, such as sitting, standing, cycling and mental arithmetic. As expected the reliability and accuracy of PPG HR and beat-to-beat detection decreases with increasing amount of arm and hand movements. Although quite significant differences between the two PPG wrist-worn devices were observed, especially with the beat-to-beat detection accuracy, the overall accuracy suggests that both of these devices can potentially be very effective in monitoring of individuals during their daily lives. In addition, the estimation of HRV was demonstrated for PPG IBI data but the results showed that HRV can be approximated from the PPG IBI data reliably only during resting conditions." | Reliable HR and Inter-beat intervals but decreased accuracy in household work conditions (hand movements) with lower percantage of correct hearbeats (only 6% as opposed to 68% when sitting) RMSSD was only accurate when the subjects were resting |  | HR - Positive IBI - Positive/Neutral RMSSD - Positive/Neutral
001_sia_wd_emp | Van Voorhees, E. E., Dennis, P. A., Watkins, L. L., Patel, T. A., Calhoun, P. S., Dennis, M. F., & Beckham, J. C. (2022). Ambulatory heart rate variability monitoring: comparisons between the empatica e4 wristband and holter electrocardiogram. Psychosomatic Medicine, 84(2), 210–214. https://doi.org/10.1097/PSY.0000000000001010 | 2018 | Psychosomatic Medicine | External | Treatment trial for posttraumatic stress disorder | 13 | M = 52 SD = 6.17 | Philips Digitrak XT Holter recorder (Philips Healthcare, Andover, Massachusetts) |  | 1 | Ambulatory 36h recording (24h starting from bedtime used for analysis) | IBI; SDNN; RMSSD; |  | "Findings reflected that the E4 device detected significantly fewer IBIs than the Holter monitor, with greater than 10% of data missing for between 40.8% and 98.4% of all intervals across the 13 participants evaluated. This high rate of missing data prevented the use of spectral analysis to evaluate the reliability of the E4, and profoundly affected the reliability of HRV as reflected in SDNN and RMSSD"(...)"While high inter-device correlations and ICCs were observed between the E4 and Holter monitors for IBIs at 1-second (median r=0.88; median ICC=0.87) and 5-minute (median r=0.94; median ICC=0.94) intervals, reliability for 5-minute RMSSD (median r= −0.09; median ICC= −0.05) and 5-minute SDNN (median r=0.48; median ICC=0.47) were poor. Agreement between the devices on 24-hour measures of HRV was satisfactory (IBI: r=0.97, ICC=0.97; RMSSD: r=0.77 IBI=0.76; SDNN: r=0.92, IBI=0.89)." | Proper agreement, but high amount of missing data, which prevented spectral analysis and negatively affected RMSSD and SDNN analysis |  | IBI - Neutral SDNN - Neutral RMSSD - Neutral
001_sia_wd_emp | Menghini, L., Gianfranchi, E., Cellini, N., Patron, E., Tagliabue, M., & Sarlo, M. (2019). Stressing the accuracy: Wrist-worn wearable sensor validation over different conditions. Psychophysiology, 56(11). https://doi.org/10.1111/psyp.13441 | 2019 | Psychophysiology | External | Healthy: demographically heterogenous sample | 50, 40 after exclusion | M = 30.2 SD = 12.74 | FlexComp Infiniti |  | 7 | seated rest; paced breathing; orthostatic; Stroop; Speech task; slow walking; keyboard typing | HR; IBI; SDNN; RMSSD; LF; HF; SCL; SCRs; nsSCRs-FREQ; nsSCRs-AMP; |  | "E4 device provides accurate mean HR measures in both static and dynamic conditions and accurate HRV measures in resting, paced breathing, and speech recovery conditions (...) but not in dynamic conditions, including speaking. Accuracy was diminished by wrist movements, cognitive and emotional stress, nonstationarity, and larger wrist circumferences. Wrist SC measures showed neither correlation nor visual resemblance with finger SC signal, suggesting that the two sites may reflect different phenomena." | HR valid for all the conditions, but HRV not accurate in dynamic conditions and while speaking. No resemblance for SC with the finger reference device, which might be due to the differences in locations. |  | HR - Positive/Neutral IBI - Neutral/Negative SDNN - Neutral/Negative RMSSD - Neutral/Negative LF - Neutral/Negative HF - Neutral/Negative SCL - Neutral/Negative SCRs- Neutral/Negative nsSCRs-FREQ - Neutral/Negative nsSCRs-AMP - Neutral/Negative
001_sia_wd_emp | van Lier, H. G., Pieterse, M. E., Garde, A., Postel, M. G., de Haan, H. A., Vollenbroek-Hutten, M. M. R., Schraagen, J. M., & Noordzij, M. L. (2020). A standardized validity assessment protocol for physiological signals from wearable technology: methodological underpinnings and an application to the e4 biosensor. Behavior Research Methods, 52(2), 607–629. https://doi.org/10.3758/s13428-019-01263-9 | 2020 | Behavior Research Methods | External | Healthy: students | 77 |  | ProComp Infiniti System with BioGraph Infiniti Software-T7500M by Thought Technology |  | 3 | Baseline Sing a Song Task Noise Task | SCL; SCR; S-AMPL; HR; SDNN; RMSSD; |  | "The results of the protocol show that the E4 wearable is valid for heart rate, RMSSD, and SD at the parameter and event levels, and for the total amplitude of skin conductance responses at the event level when studying strong sustained stressors" | EDA only valid on event level, authors suggest it might be due to the wrist's low sensitivity to smaller stressors |  | Signal level: EDA - negative PPG - NA Parameter level: SCL - Negative SCR - Negative S-AMPL - Negative HR- Positive SD HR - Positive RMSSD- Positive Event level: SCL - Positive/Neutral SCR - Positive/Neutral S-AMPL - Positive/Neutral HR- Positive SDNN - Positive RMSSD- Positive
001_sia_wd_emp | Milstein, N., & Gordon, I. (2020). Validating measures of electrodermal activity and heart rate variability derived from the empatica e4 utilized in research settings that involve interactive dyadic states. Frontiers in Behavioral Neuroscience, 14, 148–148. https://doi.org/10.3389/fnbeh.2020.00148 | 2020 | Frontiers in Behavioral Neuroscience | External | Healthy: students | 30 | M = 26.28 SD = 6.05 | Mindware Mobile Impedance Cardiograph Device |  | 2 | Seated rest; Social conversation | HR; IBI; RMSSD; SDNN; HF; SCL; |  | "We found that the E4 wristband is a reliable tool for acquiring mean IBI and mean HR data. Further, after manual editing, the IBI time-series (...) was similar to the IBI series derived from the ECG data obtained by the MW mobile device (...). Medium to high correlations were found between the devices with respect to several HRV measures, with higher correlations during rest compared to conversation. The E4 failed to produce reliable EDA data. | HR and IBI time series valid under rest and during a naturalistic conversation (both seated) with HRV measures having lower correlations during conversation and unreliable EDA overall |  | HR - Positive IBI - Positive/Neutral RMSSD - Positive/Neutral SDNN - Positive/Neutral HF - Positive/Neutral SCL - Negative
001_sia_wd_emp | Schuurmans, A. A. T., de Looff, P., Nijhof, K. S., Rosada, C., Scholte, R. H. J., Popma, A., & Otten, R. (2020). Validity of the Empatica E4 Wristband to Measure Heart Rate Variability (HRV) Parameters: a Comparison to Electrocardiography (ECG). Journal of medical systems, 44(11), 190. https://doi.org/10.1007/s10916-020-01648-w | 2020 | Journal of Medical Systems | External | Clinical: traumatized adolescents in residential care | 15 (repeated measurements, 345 recording segments) | M = 14.46; SD = 2.4 | VU-AMS |  | 2 | Baseline (watching a video); Multiple meditation sessions (8 weeks) | IBI; HR; SDNN; RMSSD; LF; HF |  | "Significant correlations were found between Empatica E4 and VU- AMS recordings for HR, SDNN, RMSSD, and HF recordings. There was a significant difference between the devices for all parameters but HR, although effect sizes were small for SDNN, LF, and HF. For all parameters but RMSSD, testing outcomes of the two devices led to the same conclusions regarding significance. The Empatica E4 wristband provides a new opportunity to measure HRV in an unobtrusive way. Results of this study indicate the potential of the Empatica E4 as a practical and valid tool for research on HR and HRV under non-movement conditions" | High correlations for HR, SDNN, RMSSD and HF with the reference device, medium but signigicant for RR, LF and LF/HF. Predictive value (rest vs meditation) good for all parameters except RMSSD |  | IBI - Positive/Neutral HR - Positive SDNN - Positive RMSSD - Negative LF - Positive/Neutral HF - Positive
001_sia_wd_emp | Bruno, E., Biondi, A., Thorpe, S., & Richardson, M. P. (2020). Patients self-mastery of wearable devices for seizure detection: A direct user-experience. Seizure, 81, 236-240. https://doi.org/10.1016/j.seizure.2020.08.023 | 2020 | Seizure | External | Clinical: epilepsy patients | 30 | M=40.9; SD=14.1 |  |  | 5 | 5 tasks: (1) fitting the device correctly; (2) switching the device on and off; (3) charging the device on a daily basis; (4) pairing the device with a phone or tablet; (5) seeking assistance. | Performance and user experience | 24h | "In a cohort of mainly “device naïve” patients with uncontrolled severe epilepsy, our study demonstrated that many, but not all, participants had overall good device self-management skills and were able to set-up and operate our system effectively.(...) Overall, the vast majority of participants managed to wear and fit the device correctly (90.0 %), to switch it on and off (70.0 %) and to charge it (63.3 %). (…) At the same time, tasks at higher risk of error were pinpointed. In particular, pairing the device with a phone or tablet, charging and swapping the device were identified as critical steps in our setting. The same results were also reflected in the TAM-FF self-administered questionnaire, highlighting that participants were aware of their performance difficulties. Future studies using similarly operating devices should implement olutions to mitigate the risk of similar errors. Alarms and notifications alerting users to reconnect a device when unpaired or reminding them when it is time to charge may be of great assistance, as well as technical improvements of the technology (automatic pairing, wireless battery charging). Our field study has also demonstrated that only half of the cohort of patients included was entirely independent in the first week of a new device use, whilst the other half needed assistance and technical support to different extents. This finding confirms the importance for many users of an easily identifiable and accessible technical support to improve device usability and to address potential malfunctions. |  | Good performance and experience when fitting and wearing the device correctly and switching it on and off. While sufficient performance was observed when charging the device, it was not experiences as effortless with some participants reporting difficulties in this area, as well as with pairing the device. Only half of the population was independant in their management of the device with other half seeking different forms of assistance. | Positive for: wearing the device correctly; switching the device on and off; and for overall experience Neutral for: charging and pairing the device; seeking assistance.
001_sia_wd_emp | Debard, G., De Witte, N., Sels, R., Mertens, M., Van Daele, T., & Bonroy, B. (2020). Making Wearable Technology Available for Mental Healthcare through an Online Platform with Stress Detection Algorithms: The Carewear Project. Journal of Sensors, 2020, Article 8846077. https://doi.org/10.1155/2020/8846077 | 2020 | Journal of Sensors | External | Healthy students | 6 | M=20.5; SD=0.8 |  |  |  |  | Qualitative interview on usability of the device (focus seems to be put on drawbacks and issues encountered while using the device) |  | "Most participants found the Empatica E4 wearable quite big. Another drawback was that it does not provide direct feedback to the user. The electrodes for the EDA sensor are large, and they press into the skin, which can get painful after wearing it for a prolonged period of time. This is especially so, given the fact that the wearable has to be worn tight to get measurements of a sufficient quality with less artefacts. This furthermore causes increased sweating under the wristband, which increases the uncomfortable feeling. We provided a manual on how to install the Empatica software on the user’s computer, but in a few cases, an error was returned during the installation. To solve these errors, an intervention from a more technically oriented person was needed. " |  | The watch has to be worn tightly for a correct measurement which increases participants' discomfort when wearing the watch for a long time. This is additionally influenced by the big size of the watch and the EDA electrodes. Support was required during installation. Lack of direct feedback into physiology was also considered a drawback in this study. | Negative for wearability, installation process, biofeedback
001_sia_wd_emp | Nasseri, M., Nurse, E., Glasstetter, M., Böttcher, S., Gregg, N. M., Laks Nandakumar, A., Joseph, B., Pal Attia, T., Viana, P. F., Bruno, E., Biondi, A., Cook, M., Worrell, G. A., Schulze-Bonhage, A., Dümpelmann, M., Freestone, D. R., Richardson, M. P., & Brinkmann, B. H. (2020). Signal quality and patient experience with wearable devices for epilepsy management. Epilepsia, 61 Suppl 1, S25-s35. https://doi.org/10.1111/epi.16527 | 2020 | Epilepsia | External | Clinical: epilepsy patients and healthy controls | 70 | Age groups: <20 y 11 (15.7%) 20-29 y 22 (31.4%) 30-39 y 15 (21.4%) 40-49 y 9 (12.9%) 50-59 y 8 (11.4%) 60-69 y 5 (7.14%) |  | Byteflies Sensor Dot, Biovotion Everion, Activinsights GENEActiv |  |  | Survey on device preference and impression of comfort (ease of use; usability; long-term comfort; comfort during sleep; willingness to use for seizure prediction) |  | "Patients who wore multiple devices were asked which device they preferred, and only responded to the survey for the preferred device. The Empatica E4 was the preferred device more often than other devices. The Biovotion device had the most positive results for both ease of manipulation and usability, followed by the two wrist-worn devices, the Empatica E4 and GENEActiv. However, all devices generally had positive responses." |  | Empatica E4 was preffered over other devices used in this study and was assessed possitively for ease of use, usability, long-term comfort, comfort during sleep. Patients also agreed they would use it for seizure prediction. | Positive for: ease of use; usability; long-term comfort; comfort during sleep; willingness to use for seizure prediction
001_sia_wd_emp | Giggins, O. M., Doyle, J., Sojan, N., Moran, O., Crabtree, D. R., Fraser, M., & Muggeridge, D. J. (2021). Accuracy of Wrist-Worn Photoplethysmography Devices at Measuring Heart Rate in the Laboratory and During Free-Living Activities. Annual International Conference of the IEEE Engineering in Medicine and Biology Society. IEEE Engineering in Medicine and Biology Society. Annual International Conference, 2021, 6970–6973. https://doi.org/10.1109/EMBC46164.2021.9629522 | 2021 | Annual International Conference of the IEEE Engineering in Medicine and Biology Society (EMBC) | External | Healthy: varied sample | 10 | M = 30.04 SD = 8 | Polar H10 chest strap |  | 4 | Seated rest; walking (4km/h); jogging (7km/h); running (10km/h) | HR; |  | The E4 device does not accurately measure HR during treadmill walking and like previous investigations, demonstrated a substantial increase in measurement error as the intensity of exercise increased [7]. Additionally, during free-living activities, the E4 device appears to produce inaccurate measurements of HR. Mean bias increases substantially as the level of activity increases, with the highest error rate observed during the running activity domain. During non-strenuous activities such as the sitting and ADL domains, the MAPE was also observed to be outside the acceptable range. Overall, it appears that the HR signal obtained from the E4 device is greatly compromised by motion artefact caused by increasing arm movement, as in previous research | HR overestimated in static conditions, but underestimated for running and was unreliable under movement and free-living conditions |  | HR - Negative
001_sia_wd_emp | Giorgi, A., Ronca, V., Vozzi, A., Sciaraffa, N., Di Florio, A., Tamborra, L., Simonetti, I., Arico, P., Di Flumeri, G., Rossi, D., & Borghini, G. (2021). Wearable Technologies for Mental Workload, Stress, and Emotional State Assessment during Working-Like Tasks: A Comparison with Laboratory Technologies. Sensors, 21(7), Article 2332. https://doi.org/10.3390/s21072332 | 2021 | Sensors | External | University sample, no details | 17, 12 after exclusion | M = 31.1 SD = 3.7 | Shimmer3 GSR+ unit; EBNeuro BeMicro |  | 3 | Baseline; N-back task; Doctor game task; Webcall task; Different difficulty levels and stress vs no stress conditions per task | HR; SCL; |  | "The results demonstrated that the parameters computed by the consumer wearable and laboratory sensors were positively and significantly correlated and exhibited the same evidences in terms of mental states discrimination" (...) "Additionally, no statistical differences were observed in terms of EBR, HR, and SCL estimation between the two technologies considered (i.e., consumer wearable and laboratory" | High convergence of HR and SCL with the reference device |  | HR - Positive SCL - Positive
001_sia_wd_emp | Stuyck, H., Dalla Costa, L., Cleeremans, A., & Van den Bussche, E. (2022). Validity of the empatica e4 wristband to estimate resting-state heart rate variability in a lab-based context. International Journal of Psychophysiology, 182, 105–118. https://doi.org/10.1016/j.ijpsycho.2022.10.003 | 2022 | International Journal of Psychophysiology | External | Healthy: students | 79, 39 after exclusion | M = 19; SD = 1.55 | Nexus-10 MKII (Mind Media BV, Herten, the Netherlands) |  | 1 | Seated position 5 minut baseline recording | PR/HR; RMSSD; SDNN; LF; HF; |  | "Our results showed that HR, RMSSD, SDNN and LF, but not HF, were validly estimated by the E4 wristband. Furthermore, the E4 wristband could validly estimate PR with recording lengths as short as 10s. RMSSD and SDNN were validly estimated using 30s or 120 s or an average of multiple short intervals (10s), while HF likely requires longer recording intervals." | From the included parameters, HF was not validly estimated |  | HR - Positive RMSSD -Positive SDNN -Positive LF -Positive HF - Negative
001_sia_wd_emp | Ravindran, K. K. G., Monica, C. D., Atzori, G., Lambert, D., Revell, V., & Dijk, D. J. (2022). Evaluating the Empatica E4 Derived Heart Rate and Heart Rate Variability Measures in Older Men and Women. Annual International Conference of the IEEE Engineering in Medicine and Biology Society. IEEE Engineering in Medicine and Biology Society. Annual International Conference, 2022, 3370–3373. https://doi.org/10.1109/EMBC48229.2022.9871559 | 2022 | Annual International Conference of the IEEE Engineering in Medicine & Biology Society (EMBC) | External | Elderly | 35 | M = 70.8; SD = 4.9 | SomnoHD polysomnography system |  | 1 | 10h sleep recording | HR; IBI; SD1: Short term variability (ms); SD2: Long term variability (ms); S: Total variability (ms); SD1UP: Short term accelerations in heart rate (ms); SD1DOWN: Short term decelerations in heart rate (ms); SDRR: Standard deviation of RR interval (ECG) or interbeat |  | "The results demonstrate that the automatic E4 heart rate measures capture long-term HRV whilst the short-term metrics are affected by missing IBIs. Our approach provides an effective way to resolve the missing IBI issue of E4 and extracts reliable heart rate measures that are concordant with ECG." (...) "We also show that there is a significant portion of the missing IBI data which is not associated with movement, but rather associated with amplitude shifts in the BVP signal in the E4. Furthermore, the E4 reliably records blood volume pulse data, but the automatically computed inter-beat interval and heart rate estimates are inaccurate and this can be attributed to the missing IBI data." | BVP data is reliable and valid for long-term heart rate variations (SD2, SDRR), but the missing IBIs negatively affect the computations of short-term HRV estimates (SD1, SD1UP and SD1DOWN). The missing IBI data were not associated with movement but amplitude changes. |  | HR - Positive IBI - Negative/Neutral SD1: Short term variability (ms) - Negative/Neutral SD2: Long term variability (ms) - Positive S: Total variability (ms); SD1UP: Short term accelerations in heart rate (ms) - Negative/Neutral SD1DOWN: Short term decelerations in heart rate (ms) - Negative/Neutral SDRR: Standard deviation of RR interval (ECG) or interbeat - Positive
001_sia_wd_emp | Duker, L. I. S., Kwon, J., Richter, M., & Pineda, R. (2023). Feasibility of wearable sensors in the NICU: Psychophysiological measures of parental stress. Early Human Development, 183, Article 105814. https://doi.org/10.1016/j.earlhumdev.2023.105814 | 2023 | Early Human Development | External | Parent-baby dyads (NICU environment) | 17 | M=29.9; SD=6.1 |  |  |  |  | Retention and adherence; acceptability |  | "Of the 17 parent-infant dyads successfully enrolled prior to the pandemic, 100 % completed the study procedures as defined by the protocol, and no participants withdrew from the study. (…) All proxy measures for acceptability indicated that the use of wireless sensors in the NICU environment was acceptable to parent participants (e.g., recruitment and enrollment: 82 % successful, retention and adherence: 100 % successful). One hundred percent of parents verbally reported the use of wireless sensors to be acceptable in the NICU environment while engaging with their infants. They were enthusiastic about the focus on parents and welcomed a non-invasive, quantitative method to examine their stress. No parents voiced concern or negative comments related to the study procedures or to wearing the E4 wristband during the study, and none requested that study procedures be stopped or that the wristband be removed. Of the 17 enrolled dyads, 41 % (n = 7) chose to continue wearing the sensor beyond the required 10-min sensory intervention, indicating the potential acceptability of wearing the sensor for a longer duration." |  | High retention and adherence, as well as high acceptability | Positive for: retention and adherence; acceptability

## 002_sia_wd_emp
device_id | study_citation | year | journal | external_internal | population | sample_size | age | reference_device | reference_device_other | events | events_included | param_studied | test_retest_time | conclusion_authors | sum_val_rel | sum_usability | final_verdict
002_sia_wd_emp

## 003_sia_wd_vri
device_id | study_citation | year | journal | external_internal | population | sample_size | age | reference_device | reference_device_other | events | events_included | param_studied | test_retest_time | conclusion_authors | sum_val_rel | sum_usability | final_verdict
003_sia_wd_vri

## 004_sia_wd_vri
device_id | study_citation | year | journal | external_internal | population | sample_size | age | reference_device | reference_device_other | events | events_included | param_studied | test_retest_time | conclusion_authors | sum_val_rel | sum_usability | final_verdict
004_sia_wd_vri

## 005_sia_wd_fit
device_id | study_citation | year | journal | external_internal | population | sample_size | age | reference_device | reference_device_other | events | events_included | param_studied | test_retest_time | conclusion_authors | sum_val_rel | sum_usability | final_verdict
005_sia_wd_fit